## Sheet1
k | v
1W1B Beamline and XAFS Experiment Station, Experiment Manual | 1W1B beamline and XAFS experiment station.pdf
Atomic local structure research and testing, experimental technology | Transmission XAFS test
Oxidation state research testing, experimental techniques | Transmission XAFS test
Bonding conditions, experimental techniques | Transmission XAFS test
Determining the atomic neighbor structure of elements present in small amounts in a sample, experimental technique | Fluorescence XAFS
1W1B beamline and XAFS experimental station experimental technology/methods, experimental technology | Transmission XAFS, Fluorescence XAFS
Fluorescence XAFS, beamline station | 1W1B-XAFS Experiment Station
Transmission XAFS, beamline station | 1W1B-XAFS Experiment Station
Transmission XAFS, for research purposes | Atomic local structure research and testing, experimental technology
Transmission XAFS, for research purposes | Oxidation state research testing, experimental techniques
Transmission XAFS, for research purposes | Bonding conditions, experimental techniques
Fluorescence XAFS, research purposes | Determine the atomic neighbor structure of elements present in small amounts in a sample
1W1B-XAFS experimental station, powder sample preparation steps (smear tape method) | Step 1: Prepare: 3M glue #810-CQ33 (19mm), A4 paper, mortar, weighing spoon, sample frame, plastic gloves, and samples to be prepared. Step 2: Take an appropriate amount of sample and grind it to less than 400 mesh. Step 3: Cut a piece of tape about 20cm long, with the pasting side facing up, and fix both ends on the A4 paper. Step 4: Spread the ground sample evenly on the tape and spread it evenly with your fingers. Shake off any remaining powder on the surface of the tape and fold the tape in half into 4 or 8 layers. Fix the processed tape on the sample frame and wait for testing. The thickness of the tape can be adjusted later based on the pre-scanned signal. After sample preparation, clean the experimental bench in time.
1W1B-XAFS experimental station, liquid sample preparation | Step 1: Prepare polyimide tape, Teflon sample holder, disposable syringe (1ml) and scissors. Step 2: Cut off an appropriate amount of polyimide tape and completely seal the window of the PTFE sample holder. Step 3: Use a disposable syringe to draw an appropriate amount of liquid sample. Step 4: Poke a small hole on the edge of the window, inject the sample into it (do not fill it full when injecting the sample, usually three-quarters is sufficient), and wipe off the remaining sample. Step 5: Take an appropriate amount of tape to seal the small hole tightly, and place the prepared liquid sample vertically. Step 6: Cover the needle of the syringe, put it into the sharps box, and tidy the experimental table.
1W1B-XAFS experimental station, powder sample tableting | Step 1: Prepare the mold, mortar, weighing spoon, weighing paper, 3M tape (#810-CQ33), etc. Clean all surfaces in the mold that may come into contact with the sample and set aside. Step 2: Transmission sample: First use the sample in the thickness software package to calculate the required sample amount. If the amount is too small, you can add an appropriate amount of BN or other reagents (low Z elements, which play a role in tableting and molding) to ensure that it can be pressed into tablets. Generally, the dosage of BN is 50-60mg, which can be pressed into tablets. Fluorescent sample: The dosage requirement is not high, as long as it can be pressed into tablets. Usually the thickness of the sheet should not exceed 1mm. Step 3: Tablet pressing Grinding: Weigh an appropriate amount of sample and use a mortar to grind the sample particles as small as possible. The particle size is 400 mesh (less than 3um) and feels like flour. At the same time, add doped BN or other reagents to grind, mix evenly, and set aside. Step 4: Assemble the mold: Use tweezers to put in the pad 1 - put in the ground sample powder (shake it to make the powder evenly distributed) - then put in the pad 2 - put in the pressure rod, and gently rotate the pressure rod. Step 5: Tablet press operation: Place the mold on the tablet press, adjust the position to ensure that the mold is in the center of the screw rod, and tighten the screw rod. Pressurizing: Swing the pressurizing handle up and down until the pressure indication increases from 0 to a certain value (unit: MPa), then stop pressurizing. Maintain pressure for 2-5 minutes. The pressure size and time are related to the sample material. It is recommended to use a mold with a diameter of 8mm or above. For specific reference: The diameter of the mold is 6mm, the total mass of the sample is about 40mg, and the pressure indicator is about 1 (not more than 3) The diameter of the mold is 8mm, the total mass of the sample is about 60mg, and the pressure indication is about 4 (not more than 6) The diameter of the mold is 10mm, the total mass of the sample is about 80mg, and the pressure indicator is about 6 (not more than 10) Pressure relief: Turn the black knob at the bottom of the tablet press to reduce the pressure indication to 0MPa. Then tighten the black knob immediately to prevent oil leakage. Step 6: Take out the mold: loosen the screw rod, take out the mold, first remove the base of the mold, then use the metal spacer to slightly tighten the screw rod, eject the sample piece, and use tape to take out the sample. Step 7: Use alcohol to wipe clean the mold, mortar and other experimental equipment, and clean the experimental table.
1W1B-XAFS experiment station, transmission XAFS experiment operation steps | 1) Optical path alignment and detection • Harness adjustment The main optical components of the beam line are slits and double crystal monochromators, and their spatial positions are crucial to the beam quality. The user is advised not to adjust the wiring harness components without authorization. •Set laser collimated beam There is a laser beam in the experimental shed, which coincides with the X-beam trace and is used as the basis for collimating the ionization chamber and the sample. Accurate. Use X-ray sensitive paper to expose X-rays at two points, and determine the direction of the laser beam through the spatial positions of the two spots. When the sample is <5mm, or there is a special sample cell, you can ask the person on duty to calibrate the laser collimated beam to avoid the spot position being different from the sample. If the deviation is too large, the laser beam direction must not be changed without permission. 2) Adjust the detuning of monochromatic light •Why should we adjust the detuning degree? The monochromatic (monoenergetic) X-beam used in the experiment contains high-order harmonic components. The high-order components generate background noise and reduce Get the quality of your data. Fine-tuning the detuning (parallelism) of the twin crystals in the monochromator can effectively suppress higher harmonics in the beam. Portion. However, excessive detuning causes too much fundamental wave intensity loss, which is undesirable. For different experiments, set up appropriate The appropriate degree of detuning is important for improving the quality of spectrum acquisition. 3) Select the working gas ratio of the ionization chamber •Why it is necessary to change the gas ratio In this transmission XAFS experimental system, two flow-type ionization chambers are used to detect the X-ray intensity before and after the sample. The size of the chamber's absorption of X-rays is related to the type of working gas and the energy of X-rays. From theoretical analysis, it can be known that the front ionization chamber absorbs When the absorption is 20% and the post-ionization chamber absorption is more than 80%, the signal-to-noise ratio of the data obtained in the experiment is the highest. Simultaneous front and rear ionization chamber work The higher the gas consistency, the better it is to eliminate the glitch caused by the fluctuation of the light source and the monochromator on the spectrum data. influence. When changing the energy range of the monochromator during the experiment, it is necessary to change the ratio of the working gas or use a different length of ionization chamber to keep the absorption rate and gas type of the front and rear ionization chambers consistent. •Operation This site is equipped with three lengths of ionization chambers: 5cm, 17cm, and 31cm, as well as pure nitrogen and pure argon. The ratio of nitrogen to argon is 75:25. Five different ratios of working gas: 50:50 and 25:75. Figure 4 shows the results of different lengths under 20% absorption. The scanning energy range applicable to the ratio of different gases in the ionization chamber. For the selection of the front ionization chamber and its working gas, please refer to this Figure. Figure 5 shows the absorption ratio and incident light energy of a 31cm long ionization chamber under five different gas ratios. The relationship curve of the post-ionization chamber gas selection can refer to this figure. 4) Sample preparation, installation and thickness adjustment •Adjust sample thickness It can be seen from theoretical calculation that when  .t is a certain value, the best signal-to-noise ratio can be obtained. The commonly used sample thickness t is obtained by pressing  and then .t- before.t=1~1.5, where after is the  value after the absorption limit, and before is the  value before the absorption limit. From the spectrum, After.t-Fore.t is absorption. The height of the limit jump edge. The actual situation is more complicated. Due to the diverse composition of various elements in the sample, some samples cannot be analyzed. The above requirements are met, and the difference does not change monotonically with the increase of t. If the sample is too thick or too thin, the jump limit height will be decrease, the optimal value can only be found through experimentation. For thickness testing, see Part 4 of the XAFS data acquisition software. Through testing, a spectrum near the absorption edge can be obtained, adjust the sample thickness according to its edge jump height value, and repeat the above steps. Until the jump edge height reaches or is closest to 0.6~1.6. •Sample preparation For transmission XAFS experiments, it is recommended to use powder samples with a powder fineness of 400#. Take a width of 1-2cm Tape (3M brand tape is recommended) about 20cm, sprinkle the powder evenly on the adhesive surface and spread evenly, fold the tape to Easily adjust thickness. This site is equipped with a tablet press, and powder samples can also be pressed into tablets for sample preparation. For flaky samples that cannot be made into powder by spinning or sintering, the thickness should be controlled below 100m and the area should be larger than 5×10mm and must not have holes. Prepare 2 to 3 pieces for thickness adjustment. For liquid state samples, it is recommended to use a sample box. 5) Settings of data acquisition electronic system •User concerns In the data acquisition electronics system, what users need to pay attention to are the two preamplifiers placed on the lower floor of the experimental shed. (KEITHLEY 428) Magnification setting and background current adjustment. In the data acquisition system, the output current of the front and rear ionization chambers is weak The signal is amplified by the preamplifier and converted into a voltage signal and sent to the V/F converter to be converted into a pulse signal proportional to the frequency. It is then counted by the scaler and fed into the computer. The input signal range of the V/F converter is 0-10V. There is a signal circuit in the control cabinet The voltage monitoring unit displays the signal voltage value output by the preamplifier in real time, and there is an audio alarm when the voltage exceeds the range. if The magnification setting is inappropriate. When the experiment changes the energy range or due to changes in signal voltage during the spectrum collection process, overshoot may occur. range, causing spectrum collection to fail. On the other hand, in the weak signal area, the system nonlinearity increases and the noise influence also becomes larger, so the amplifier The output signal of the amplifier should not be below the 0.01V level. Based on the above reasons, before officially collecting spectrum, set the preamplifier appropriately. The amplifier magnification is a very critical step. •Set up the preamplifier While running the thickness test in the XAFS data acquisition software package to adjust the sample thickness, attention should be paid to the According to the display of the signal voltage monitoring unit, during the entire thickness test scanning process, the maximum value of the two signals should be between 0.8-9V. time, if the value is less than this value, the amplification factor of the preamplifier should be increased, otherwise it should be decreased. 6) Collect music scores For the operation of the spectrum acquisition software, please refer to the Beijing Synchrotron Radiation XAFS Experiment Station Spectrum Acquisition Program Instructions (Appendix 1) for experimental parameters. Here are some additions to the settings section. •Principles for selecting absorbing edges The photon energy scanning range provided by this experimental station is 4.8~22.8keV (due to the absorption of low-energy X-rays in the air To collect heavier, low-concentration Ti/Cs samples, it is recommended to go to Zhongneng Station 4B7A). Generally, the spectral range of XAFS experiments is between Around 1keV, when selecting the absorption edge of the test element, the K edge is first considered. If the spectrum range exceeds 4.8~22.8keV, Or if the spectrum range overlaps with the absorption edge of other elements in the sample, the L edge must be considered. Among L sides, select at least the following number of sides There is no overlap within the range of hundreds of eV. According to this principle, the LIII side is generally taken. •Set sampling time The longer the sampling time, the smaller the statistical error of spectrum collection, but the entire spectrum collection time is lengthened. Light intensity and sampling time are affected by There are two factors that affect the statistical error of the data. Based on the light intensity of this beamline, a sampling time of 1 to 2 seconds is generally available in transmission mode.
1W1B-XAFS experimental station, research area | Materials Science
1W1B-XAFS experimental station, research area | environmental science
1W1B-XAFS experimental station, research area | nanoscience
1W1B-XAFS experimental station, research area | biomedicine
1W1B-XAFS experimental station, research area | Chemistry and Chemical Engineering
1W1B-XAFS experimental station, research area | Energy Catalysis
1W1B-XAFS experimental station, research area | Humanities and Archeology
Main parameters and equipment of 1W1B-XAFS beam line and experimental station | Energy range available for experiments: 4.8 — 22.8 keV Energy resolution: <1-3 × 10-4 @ 9 keV Luminous flux at sample: >1×1011photons/s @ 10 keV; Isr = 250 mA@2.5GeV Spot size at sample (FWHM): 1(H) ×0.5(V) mm2 Flow type ionization chamber (detector): Gas: pure nitrogen, pure argon, nitrogen-argon mixture Ionization chamber length: 31cm/17cm/5cm Working voltage: 2200V/2200V/1000V Fluorescence ionization chamber (detector): Lyttle Fluorescence Ionization Chamber Detector 19 yuan high purity germanium semiconductor array detector Low temperature sample room Temperature control range: 10 - 300 K
1W1B-XAFS experimental station, line station status | normal operation
1W1B beamline and XAFS experimental station, Beijing synchrotron radiation experiment application process | Step 1: Project application, ● Ordinary project application: Users register and log in to the Chinese Academy of Sciences Major Science and Technology Infrastructure Sharing Service Platform (http://lssf.cas.cn/) (hereinafter referred to as the platform) to submit a project application, then download and print the application form, sign it at the designated location and stamp it with the official seal of the unit, and mail one copy to the BSRF User Office (hereinafter referred to as the user office). The User Office accepts project applications throughout the year. ● Application for key topics: Key topics must meet the following conditions: basic research or applied basic research with significant academic value; applied research that is expected to achieve significant economic benefits or social impact in the short term (usually three years); especially research projects that have received major or key national support, or research that is current research hotspots and frontiers; applicants have good research performance in related fields, have experience in synchrotron radiation work, and have a stable research team. Users who have carried out experiments at BSRF for more than two years can apply for installation key projects. Users who want to apply for key topics also need to submit an application online through the platform, and mail a copy of the application signed and stamped with the official seal of the unit to the user office. The specific reception time for key project applications is subject to the notice issued by the user office, which is generally once a year. Step 2: Project review● Preliminary review of the experimental station: evaluate the feasibility of the experimental technology;● Two peer experts review: scientific significance and academic importance (sent twice for review in March and September each year);● After passing the review of general topics, project management (valid for two years); key topics also need to undergo on-site report defense. After being reviewed and approved by the professional committee and user committee, project management (valid for two years) will be given priority. Step 3: Apply for opportunities and arrange opportunities. ● Initial review of the experimental station: assess the feasibility of the experimental technology; ● Review by two peer experts: scientific significance and academic importance (submitted twice for review in March and September each year); ● After passing the review of general topics, project management (valid for two years); key topics require on-site report defense. After review and approval by the professional committee and user committee, project management (valid for two years) will be given priority. Step 4: Complete safety training and conduct experiments● First-time experimenters need to register an account on the platform and complete personal radiation protection safety training. Only after passing the exam can they enter the experimental hall (the safety training is valid for 2 years);● After entering the experimental line station, users must strictly Comply with the various rules and regulations of BSRF; fill in the light consumption record form, conduct online station safety operation training under the guidance of online station personnel, and then carry out experiments; ● After the experiment, users need to clean the experimental bench and dispose of experimental waste in accordance with relevant regulations; return the access card and personal dose card. Step 5: Conclusion and submission of results ● When publishing user scientific and technological achievements, if all or part of the experimental content is completed on BSRF, it must be noted that "data were obtained at Beijing Synchrotron Radiation Facility XX Station", and relevant achievement information (including articles, awards, patents, social benefits, etc.) should be fed back to the user office in a timely manner.
1W1A-Diffuse Scattering Experimental Station, Research Area | crystalline material thin film
1W1A-Diffuse Scattering Experimental Station, Research Area | Multilayer film nanomaterials
1W1A-Diffuse scattering experimental station, research content | The X-ray diffuse scattering experimental station is mainly used to carry out structural research on the distribution and configuration of various defects in crystal materials (vacancies, impurities or interstitial atoms, dislocations, dislocation loops, stacking faults, atomic groups or vacancies, etc.), surface and interface research on thin films and multilayer film structures, high-resolution X-ray diffraction research, and exploratory research on new experimental techniques and methods.
1W1A-Diffuse Scattering Experimental Station, Experimental Technology | X-ray reflectivity (XRR)
1W1A-Diffuse Scattering Experimental Station, Experimental Technology | X-ray non-specular scattering
1W1A-Diffuse Scattering Experimental Station, Experimental Technology | Grazing incidence X-ray diffraction (GIXRD)
1W1A-Diffuse Scattering Experimental Station, Experimental Technology | Grazing incidence X-ray scattering (GIXS)
1W1A-Diffuse Scattering Experimental Station, Experimental Technology | Grazing incidence small-angle X-ray scattering (GISAXS)
1W1A-Diffuse Scattering Experimental Station, Experimental Technology | Grazing incidence wide-angle X-ray scattering (GIWAXS)
X-ray reflectivity (XRR), research content | Analyze film thickness, roughness, density
X-ray non-specular scattering, research content | Thickness, density, surface roughness, and interface correlation length of the film sample
Grazing incidence X-ray diffraction (GIXRD), research content | Diffraction structure information of test samples
Grazing incidence X-ray scattering (GIXS), research content | Near-surface structure of thin film materials
Grazing incidence small-angle X-ray scattering (GISAXS), research content | ① Measurement of surface and surface structure information under normal pressure. The main materials studied include surface nanoparticles, surface-embedded nanometal particles, mesoporous materials, block copolymer films, multilayer films, quantum dots, quantum wires, quantum wells, etc. in the semiconductor nanofield. ② Mainly in-situ experiments in ultra-high vacuum environment, mainly including the growth of metal oxide material interfaces and surface nanostructures, the catalytic process of nanocatalytic particles, the self-assembly process of nanoparticles and copolymers, etc.
Grazing incidence wide-angle X-ray scattering (GIWAXS), research content | Diffraction structure information of test samples
1W1A-Diffuse scattering experimental station, GIWAXS batch test sample preparation/preparation | 1 Size and material of sample Substrate size: It is recommended to be 10-20mm square and the thickness should not exceed 2mm. Substrate material: silicon wafer, quartz wafer, glass wafer can be used, the surface is required to be smooth, flat and hard. 2 Sample number Number all samples in advance, starting from the number 1, and write the numbers on the blank space of the plexiglass plate. 3 Pasting of samples Paste approximately 5-10mm square double-sided tape on the center of the back of the sample substrate and stick it on the sent plexiglass plate. Note: Please stick the sample on the correct side of the plexiglass plate as shown in the picture below (please pay attention to the arrangement direction of the three magnets, the current sticking side is opposite to the previous old version). Please place the three magnets as close to the center of the sample piece as possible. X-rays will be incident in the direction shown by the arrow in the figure. Note: The sample and the plexiglass plate must be firmly and tightly pasted! There must be no double-sided tape on the back of the sample holder, otherwise it may be adsorbed on the gripper of the manipulator. If the size of the sample is small, pay attention to the sticking position of the sample and be sure to cover the three magnets. If the size of the sample is not enough to cover the three magnets, the sample loading will fail. ; 4 Preparation of organic glass plate for pasting samples Remove the protective film on both sides of the plexiglass plate. Use transparent tape to stick one side of the plexiglass plate, and place this side on another plexiglass plate with magnets attached (a standard plate with magnets will be mailed). Place three magnets so that they can stick to the existing magnets below (if they cannot stick, put them in reverse). Wrap the magnet with tape and wrap it tightly twice (one circle is not enough, make sure it is tightly adhered). ; 5 Storage or mailing of batch samples An iron box (such as a pencil box) can be used, and the organic glass plate with the sample stuck on it can be sucked into the box. Note: There is no need to stick double-sided tape between the plexiglass plate and the iron box (the magnets on the plexiglass plate can already be firmly attached), otherwise it will affect the picking and placing, and the remaining double-sided tape may stick to the gripper of the robot. If express delivery is required, a soft protective layer can be added and the height of the iron box should be as short as possible.
1W1A - Diffuse Scattering Experiment Station, GIWAXS Batch Test Sample Preparation/Preparation Reference Connection | https://mp.weixin.qq.com/s/vujYkQMcIzOcaIMYjZ5Bqg
1W1A-Diffuse scattering experimental station, equipment parameters | 1W1A beamline performance parameters: Monochromator: Dimensions (mm3) --- 100(l) × 30(w) × 1(t) Bevel angle (°) --- 14.7 Diffraction angle (°) --- 23.65 Crystal bending radius (m) --- 31 Distance from light source to crystal (m) --- 20 Crystal to sample distance (m) --- 5 Reflector: Dimensions (mm2) --- 800(l) × 80(w) Grazing incidence angle (°) --- 0.2 - 0.35 Mirror bending rate radius (m) --- 1740.87 - 994.78 Distance from light source to mirror (m) --- 21.6 Mirror to sample distance (m) --- 3.4 Sample point: Focused spot size (mm2) --- 0.7(l) × 0.4(w) Monochromatic light energy resolution (Si(220)) --- < 4.4 × 10-4 Second harmonic (Si(220)) --- < 0.2%
1W1A-Diffuse Scattering Experiment Station, main equipment of X-ray diffuse scattering experiment station | Five-circle diffractometer
1W1A-Diffuse Scattering Experiment Station, main equipment of X-ray diffuse scattering experiment station | detector system
Five-circle diffractometer, equipment introduction | A Huber four-circle diffractometer that can flip around the incident light in horizontal and vertical positions, with a minimum step size of 9" for the theta circle and a minimum step size of 9" for the 2theta circle, an absolute rotational accuracy of 30", and a repeatability accuracy of 2" (Figure 3 is a photo of the five-circle diffractometer in grazing incidence).
Detector system, equipment introduction | NaI scintillation counter.
1W1A-Diffuse scattering experimental station, ordinary XRD experimental measurement | After cutting the sample into the optical path in a similar sequence to the reflection measurement, the next step is to find the diffraction peak, which can be performed as follows (taking the single crystal Si wafer (004) sample as an example): Adjustment purpose: θ-2θ is placed at the diffraction peak of Si (004) Adjustment method: 1: Run the 2θ angle to 69.1 degrees, open the 2θ detector slit as large as possible, and add an absorption sheet to prevent detector saturation; 2: Do a θ scan, and an obvious diffraction peak can be observed near θ A. Then scan θ carefully near peak A to determine the peak A*, and place θ at the peak A*; 3: Scan the χ angle within a certain angle range to find the strongest diffraction position χ0, and place the sample at this position. 4: Set the 2θ detector slit to a suitable size, generally 0.2~0.5mm, add an appropriate absorption sheet in front of the detector, do a 2θ fine scan near 69 degrees, determine the peak B*, run the detector to the peak position B*, set 2θ to 69.1 degrees, and the adjustment is completed. [Note] Be sure to add an absorbing sheet when the 2θ detector slit is opened. After the above adjustment is completed, the incident X-rays, Si (004) crystal plane and the position of the detector meet the Bragg diffraction conditions. After the adjustment is completed, the user can conduct experiments according to their own experimental purposes.
1W1A-Diffuse scattering experimental station, grazing incidence diffraction (GID) measurement | Grazing Incidence Diffraction (GID) Measurements I. Determine the 2θ detector and post-detector zero points; [Explanation] The 2θ detector is used for experimental measurements, and the rear detector is used for cut-light and grazing incidence angle calibration. 1. Remove the sample stage, set the 2θ detector slit to 0.2mm, and set the rear detector slit to 0.2mm 2. Add an appropriate amount of absorbent tablets (usually 12 to 15 tablets) [Note] It is strictly forbidden to place the detector in a position aligned with direct light without adding an absorbing sheet. 3. Turn on the light and scan with the 2θ detector For example: ascan tth –1 1 200 0.1 If the 2θ peak position is found at CEN through scanning, place the 2θ detector at this position and set this position as the zero point. The operation is as follows: mv tth CEN (Place the 2θ detector at the peak position just scanned) set tth 0 (set the peak position to the 2θ zero point coordinate) 4. Move the 2θ detector away and select the mapping detector as the rear detector. mv tth 60 (remove 2θ detector) plotselect (plot selection post-detector) Select bdet ascan psdy -0.2 0.2 40 0.1 in the prompt options Just like setting the zero point of the 2θ detector, set the coordinates of the zero point of the rear detector. mv psdy CEN set psdy 0 If the peak cannot be found during the scan at this time, the scan range can be expanded. Then ct , record the reading and use it for light cutting [Tip] Two commonly used commands: ct (record readings from individual detectors and ionization chambers) wa (used to observe each motor, that is, the position of the detector) II. Place the sample 1. Place the sample. Place the sample stage on the instrument 2. Move the 2θ detector to a 60-degree position to prevent light blocking mv tth 60 3. Move chi to -9 degrees: mv chi –9 (experience value) 4. Turn on the laser, disengage the motor of the Ф circle (sample circle), record the initial reflection spot position, then rotate the sample 180 degrees, record the spot position again, and adjust the sample elevation so that the spot is located at the center of the line connecting the two positions. Then rotate 180 degrees in the opposite direction and adjust repeatedly until the reflected light spot does not move significantly when the sample is rotated. During the rotation process, pay attention to the connection of the sample lifting motor on the back. Do not continuously rotate the sample in the same direction, otherwise the motor connection will be entangled and cause an accident. 5. Close the motor of the Ф circle (sample circle) and turn off the laser. III. Cut light [Explanation] The purpose of light cutting is to make the sample surface parallel to the incident light beam, and to position the sample in a position that blocks half of the incident light in the vertical direction. The light cutting is performed using a rear detector. Note that the reading of the rear detector is much smaller than that of the front detector and cannot exceed 200,000/second. 1. Use the rear detector and move it to the zero point: mv psdy 0 2. Lower zz to let all the light pass through, and record the count of the ionization chamber at this time umvrzz-1 ct 3. Gradually raise the zz sample stage and count mvrzz 0.1 ct Repeat the above operation until the light intensity is reduced to approximately half 4. The sample is scanned in elevation to make the surface of the sample parallel to the beam; ascan th –2 2 100 0.1 mv th CEN set th 0 5. Cut the light again and place zz in a position that blocks half of the beam, that is, the count of the ionization chamber is reduced to half of the count when the entire beam passes. At this time, you can gradually move zz upward with a very small step size, and always pay attention to whether the count of the ionization chamber has been reduced to half: mvzz 0.01 ct IV. Determination of reflection angle The principle of determining the reflection angle is to use the reflected beam to determine the incident angle: 1. Depending on the reflectivity of the sample, keep 2 to 3 absorbing sheets or remove the absorbing sheets completely. 2. Determine the angle of incidence mvpsdy 0.4 ascan chi 0 0.4 40 0.1 Find peak position CEN mvchi CEN set chi 0.2 V. Diffraction angle determined [Explanation] The substrate is generally used to determine the diffraction angle position. For convenience of explanation, this explanation uses Si (400) diffraction surface as an example 1. [Note] Remove the front absorption sheet; place a thick absorption aluminum sheet to block the rear detector; fully open the 2θ detector slit; 2. Select the mapping detector as the 2θ detector: plotselect Select det in the options; 3. Place the 2θ detector at the Si (400) crystal plane peak position: 69.1 degrees, and scan the ϕ angle For example ascan phi 0 90 450 0.1 In order to find the peak position, you can quickly scan a large range, roughly determine the peak position, and then perform a fine ϕ scan to determine the θ diffraction peak position. 4. Place ϕ at the maximum peak position, such as mv phi CEN, then reduce the 2θ detector slit to 0.3~0.1mm, and perform 2θ fine scanning. 5. Place the 2θ detector at the peak position of the fine scan, and then finely scan the ϕ angle of the sample. Set the above 2θ detector position to 69.1 degrees and the sample ϕ angle to 34.55 degrees; 6. If you think that the above ϕ angle peak strength is not enough, you can increase the ϕ angle of the sample by 90 degrees or 180 degrees and repeat steps 4 and 5 for measurement; After calibration, you can start testing according to your needs.
1W1A-Diffuse scattering experimental station, precautions | [Notes]: 1. The front detector reading must be less than 1 million/second to prevent oversaturation damage; usually, when the detector reading is greater than 500,000, it is necessary to observe and analyze whether it is oversaturated; the rear detector reading must be less than 200,000/second, and when it exceeds 100,000, attention must be paid. 2. It is strictly prohibited to place each detector in a direct light position without adding an absorbing sheet to prevent saturation damage; 3. The minimum step size of ϕ is 0.005, and the minimum step size of 2θ is 0.0025. Pay attention to the step size setting to prevent it from being lower than the minimum step size. To save time, you can increase the step size and reduce the number of steps. 4. The file name of the data cannot be the word "log". 5. How to use a U disk to copy data: Click the start menu in the lower left corner of the computer (the image is a little red hat), and navigate to: System toolsÆhardware browserÆhard drives. The following image will appear: Pay attention to the three characters "sdd" in the last line of the picture above. Different U disks have different corresponding characters. If you are using a USB flash drive If it is not recognized by the system, there will be no display of this line of characters, as shown in the following figure: Open a new command terminal and enter in the command line: Mount –t vfat /dev/sdd /mnt/usb, then U When mounting the USB flash drive, please note that the three characters "sdd" will change with different USB flash drives; after copying the experimental data, enter in the command line when unmounting the USB flash drive: Umnout /mnt/usb
1W1A-Diffuse scattering experimental station, line station status | normal operation
1W1A-Diffuse Scattering Experiment Station, Beijing Synchrotron Radiation Experiment Application Process | Step 1: Project application, ● Ordinary project application: Users register and log in to the Chinese Academy of Sciences Major Science and Technology Infrastructure Sharing Service Platform (http://lssf.cas.cn/) (hereinafter referred to as the platform) to submit a project application, then download and print the application form, sign it at the designated location and stamp it with the official seal of the unit, and mail one copy to the BSRF User Office (hereinafter referred to as the user office). The User Office accepts project applications throughout the year. ● Application for key topics: Key topics must meet the following conditions: basic research or applied basic research with significant academic value; applied research that is expected to achieve significant economic benefits or social impact in the short term (usually three years); especially research projects that have received major or key national support, or research that is current research hotspots and frontiers; applicants have good research performance in related fields, have experience in synchrotron radiation work, and have a stable research team. Users who have carried out experiments at BSRF for more than two years can apply for installation key projects. Users who want to apply for key topics also need to submit an application online through the platform, and mail a copy of the application signed and stamped with the official seal of the unit to the user office. The specific reception time for key project applications is subject to the notice issued by the user office, which is generally once a year. Step 2: Project review● Preliminary review of the experimental station: evaluate the feasibility of the experimental technology;● Two peer experts review: scientific significance and academic importance (sent twice for review in March and September each year);● After passing the review of general topics, project management (valid for two years); key topics also need to undergo on-site report defense. After being reviewed and approved by the professional committee and user committee, project management (valid for two years) will be given priority. Step 3: Apply for opportunities and arrange opportunities. ● Initial review of the experimental station: assess the feasibility of the experimental technology; ● Review by two peer experts: scientific significance and academic importance (submitted twice for review in March and September each year); ● After passing the review of general topics, project management (valid for two years); key topics require on-site report defense. After review and approval by the professional committee and user committee, project management (valid for two years) will be given priority. Step 4: Complete safety training and conduct experiments● First-time experimenters need to register an account on the platform and complete personal radiation protection safety training. Only after passing the exam can they enter the experimental hall (the safety training is valid for 2 years);● After entering the experimental line station, users must strictly Comply with the various rules and regulations of BSRF; fill in the light consumption record form, conduct online station safety operation training under the guidance of online station personnel, and then carry out experiments; ● After the experiment, users need to clean the experimental bench and dispose of experimental waste in accordance with relevant regulations; return the access card and personal dose card. Step 5: Conclusion and submission of results ● When publishing user scientific and technological achievements, if all or part of the experimental content is completed on BSRF, it must be noted that "data were obtained at Beijing Synchrotron Radiation Facility XX Station", and relevant achievement information (including articles, awards, patents, social benefits, etc.) should be fed back to the user office in a timely manner.
1W2B-Comprehensive Experiment Station, Experiment Manual | Development of multi-method experimental techniques, conventional absorption spectroscopy, and methodological research based on high-performance detectors
1W2B-Comprehensive Experiment Station, Introduction | BSRF 1W2B - Comprehensive Experiment Station is located in the North Hall of Synchrotron Radiation (Hall 15). It has been transformed from a dedicated crystallography line station into a comprehensive experimental station. It mainly carries out conventional hard X-ray absorption spectroscopy experiments, laser pump-X-ray probe time-resolved experiments, SAXS/WAXS/XAFS joint measurement experiments and XRD/XAFS joint measurement experiments. The line station provides continuously adjustable monochromatic X-ray from 5-18keV. Absorption spectroscopy experiments can be carried out in two modes: transmission and fluorescence. The fluorescence method uses Lytle detectors or self-developed array SDD systems. Time-resolved experiments can provide time-resolved XAFS experiments or XRD experiments of laser pump-X-ray probes, with a time resolution of up to 120 picoseconds. The joint measurement experiment can simultaneously collect the absorption spectrum, small-angle and wide-angle signals of the sample or the diffraction and absorption spectrum signals of the sample simultaneously. (Diffraction and wide-angle signal collection use a Mythen detector array with one-dimensional position resolution and no two-dimensional position resolution capability.)
1W2B-Comprehensive Experiment Station, Beijing Synchrotron Radiation Experiment Application Process | Step 1: Project application, ● Ordinary project application: Users register and log in to the Chinese Academy of Sciences Major Science and Technology Infrastructure Sharing Service Platform (http://lssf.cas.cn/) (hereinafter referred to as the platform) to submit a project application, then download and print the application form, sign it at the designated location and stamp it with the official seal of the unit, and mail one copy to the BSRF User Office (hereinafter referred to as the user office). The User Office accepts project applications throughout the year. ● Application for key topics: Key topics must meet the following conditions: basic research or applied basic research with significant academic value; applied research that is expected to achieve significant economic benefits or social impact in the short term (usually three years); especially research projects that have received major or key national support, or research that is current research hotspots and frontiers; applicants have good research performance in related fields, have experience in synchrotron radiation work, and have a stable research team. Users who have carried out experiments at BSRF for more than two years can apply for installation key projects. Users who want to apply for key topics also need to submit an application online through the platform, and mail a copy of the application signed and stamped with the official seal of the unit to the user office. The specific reception time for key project applications is subject to the notice issued by the user office, which is generally once a year. Step 2: Project review● Preliminary review of the experimental station: evaluate the feasibility of the experimental technology;● Two peer experts review: scientific significance and academic importance (sent twice for review in March and September each year);● After passing the review of general topics, project management (valid for two years); key topics also need to undergo on-site report defense. After being reviewed and approved by the professional committee and user committee, project management (valid for two years) will be given priority. Step 3: Apply for opportunities and arrange opportunities. ● Initial review of the experimental station: assess the feasibility of the experimental technology; ● Review by two peer experts: scientific significance and academic importance (submitted twice for review in March and September each year); ● After passing the review of general topics, project management (valid for two years); key topics require on-site report defense. After review and approval by the professional committee and user committee, project management (valid for two years) will be given priority. Step 4: Complete safety training and conduct experiments● First-time experimenters need to register an account on the platform and complete personal radiation protection safety training. Only after passing the exam can they enter the experimental hall (the safety training is valid for 2 years);● After entering the experimental line station, users must strictly Comply with the various rules and regulations of BSRF; fill in the light consumption record form, conduct online station safety operation training under the guidance of online station personnel, and then carry out experiments; ● After the experiment, users need to clean the experimental bench and dispose of experimental waste in accordance with relevant regulations; return the access card and personal dose card. Step 5: Conclusion and submission of results ● When publishing user scientific and technological achievements, if all or part of the experimental content is completed on BSRF, it must be noted that "data were obtained at Beijing Synchrotron Radiation Facility XX Station", and relevant achievement information (including articles, awards, patents, social benefits, etc.) should be fed back to the user office in a timely manner.
1B3-X-ray lithography experimental station, experimental manual | 1B3A X-ray Lithography Station.pdf
1B3-X-ray lithography experimental station, introduction | The X-ray lithography station adjusts the light path through a reflector and can achieve hard line exposure (about 4.3keV) and soft line exposure (about 1.6keV) respectively. Hard line exposure uses synchrotron radiation white light to reach the sample through a beryllium window with a thickness of about 250um. It can achieve exposure of PMMA with a thickness exceeding 1.5mm and is mainly used in the development of microstructures with large aspect ratios. In soft line exposure, white light is reflected by a chromium mirror and then reaches the sample through a silicon nitride window with a thickness of about 0.5um. It can achieve photolithography of lines with a width of about 100 nanometers. It is mainly used in the development of sub-micron microstructures and mask preparation for hard line exposure.
1B3-X-ray lithography experimental station, research area | Development of X-ray optical components, such as gratings, refractive lenses, etc.
1B3-X-ray lithography experimental station, experimental technology | nanolithography
1B3-X-ray lithography experimental station, experimental technology | Depth X-ray Lithography
Nanolithography, research content | nanolithography
Deep X-ray lithography, research content | Metal gratings, refractive focusing lenses, zone plates, precision micro-metal structure systems
1B3-X-ray lithography experimental station, research content | It is mainly used for the preparation of X-ray optical components, especially for components containing large aspect ratio microstructures, such as X-ray refractive lenses, zone plates, gratings, etc. The achieved microstructure size and thickness cover a wide range, and its devices are widely used in the fields of X-ray focusing and imaging.
1B3-X-ray lithography experimental station, main equipment | Main equipment of the experimental station In addition to the exposure system of the synchrotron radiation lithography beam line, the lithography experimental station also has relatively complete micro-processing equipment and environment. The working environment includes a 300-square-meter thousand-level clean room, corrosion room, process auxiliary room, etc.; the equipment has a magnetron sputtering system, an atomic layer deposition system, a thermal evaporation system, an electroplating device, an oxidation furnace, a UV lithography machine, a reactive ion etching machine, an ion beam etching machine, and an ICP etching system; and testing equipment includes a step meter, a microscope, etc.
1B3-X-ray lithography experimental station, light source performance and parameters | The X-ray lithography station (1B3A) of Beijing Synchrotron Radiation Facility is led by 1B3 bent iron. By adjusting the mirror system, two optical paths of reflected light and direct light can be realized, respectively, for nano-lithography suitable for sub-micron line width and X-ray depth lithography suitable for thickness greater than 1 mm. The reflector can move vertically to switch the light path. When the reflector moves away from the light path, the synchronization light enters the X-ray depth lithography exposure cavity; when the reflector cuts into the light path, the synchronization light enters the nanolithography exposure cavity after being reflected by the reflector, so the two exposure modes must be used in a time-sharing manner. The synchrotron radiation beam is uniformly distributed in the horizontal direction and Gaussian in the vertical direction. Therefore, the sample stage is scanned in the vertical direction to meet the large-area uniform spot required for photolithography. Figure 1 is a schematic structural diagram of the 1B3A photolithography beamline.
X-ray depth lithography, equipment parameters | X-ray depth lithography The 1B3A beam line is led from a bent iron, and the light source is illuminated on the sample through a beryllium window with a thickness of 250 microns. The spot size is approximately 55 mm horizontally and 10 mm vertically. The sample environment is a 50kpa helium atmosphere. The sample stage scans and exposes in the vertical direction. The maximum travel range is ±50 mm. Fixed-point exposure without scanning can also be achieved. Wavelength range: White light, peak value approximately 4.3kev. Exposure area: Maximum 50*80 mm. Lithography Depth: Maximum lithography depth 1.5mm for PMMA photoresist.
Nanolithography, equipment parameters, | nanolithography The 1B3A beamline light source is reflected by a chromium-plated reflector, then passes through a 0.5-micron-thick silicon nitride film, and then passes through an 18-micron-thick beryllium film (optional) before illuminating the sample. The spot size is approximately 50 mm horizontally and 10 mm vertically. The sample environment is a vacuum, and the sample stage scans and exposes in the vertical direction, with a maximum travel range of ±50 mm. Fixed-point exposure without scanning can also be achieved. Wavelength range: White light, peak value approximately 1.6kev. Exposure area: 50*80 mm. Lithography resolution: minimum line width 50nm.
1B3-X-ray lithography experimental station, X-ray exposure operation equipment | 1 Exposure chamber, water cooler, gas cylinder Hard line exposure chamber: used for placement and scanning of hard line samples. Soft wire exposure chamber: used for placing and scanning soft wire samples. Hard line water cooler: used for cooling samples during hard line exposure. Cord water cooler: used to cool samples during cord exposure. Helium bottle: Fills the hardline exposure chamber with helium. 2 Safety interlock control Hutch door control box: used for safety interlock control of Hutch doors. Hutch door control box: used for safety interlock control and ventilation and lighting switches of Hutch doors. Safety interlock control system: used for switch control of exposure time. 3 Exposure control Exposure control screen: mainly used for the operation of exposure control software, and also used for line stations Equipment monitoring, etc. Exposure control software: used for setting, changing, collecting, etc. exposure. 4 Vacuum control operation cabinet Touch screen: Sample scanning control software operation for hard/soft lines. Hard line scan control software: used for scan control of hard line samples. Soft line scanning control software: used for scanning control of soft line samples. Hard wire vacuum control: control vacuuming, helium filling, deflation, etc., and monitor vacuum and helium filling. Soft line vacuum control: used for vacuuming, switching valves, deflating, etc. Soft line vacuum gauge: used for soft line vacuum monitoring. Power supply area: the total power supply of the entire vacuum control cabinet and the independent power supply of each part.
1B3-X-ray lithography experimental station, hard line exposure operation steps | Hard line exposure operation steps (1) Turn on the cooling water. Open the plug button on the hard-wired water cooler and let the external circulation pump start. Turn on "POWER" on the water cooler panel and press the "COOL" button after about one minute. (2) Place the sample. Fix the sample in the sample holder, insert the sample holder into the scanning frame in the hard-line exposure chamber, and tighten the screws to secure it. Close the exposure chamber door. (3) Start scanning. On the touch screen hard-line exposure control system, set the "speed setting" and "swing setting" according to your own experimental needs. (Note that the parameters after "center height" cannot be changed.) After the settings are completed, click "Scan Start". (4) Hutch door interlocking. Press the square green "Search" button inside the Hutch door, and the door warning light will flash and beep. Press the "Close" button outside the Hutch door and the door will automatically close. After the door is closed, click the square green "Search" button outside the door. After the beep ends, the round green light (Enable) under the safety gate on the safety interlock control system panel will light up. (5) Vacuum. On the hard-wired control power box, press the "on" button of the mechanical pump, and the mechanical pump will start. Then press the "on" button of the front-stage valve, and you can hear a loud pumping sound, which will gradually become smaller. Evacuate to about 8Pa, press the "on" button of the foreline valve to close the foreline valve, and then press the "off" button of the mechanical pump. (6) Fill with helium. Press the "Open" button of the helium filling valve and rotate the flow control valve on the helium bottle counterclockwise to fill the exposure chamber with helium. When the pressure reaches about 50KPa and the pressure difference is less than 1KPa, the helium filling valve will automatically close. Turn the flow control valve on the helium bottle clockwise to close. (7) Open the gate valve. Press the "open" button of the gate valve to open the gate valve. (8) Set exposure. On the LIGA light intensity acquisition software interface, enter the exposure amount in the box below "Target current value mAH", then click on the right-facing hollow arrow in the upper left corner, which changes to a solid arrow after clicking. (9) Start exposure. Click the "Start" button on the LIGA light intensity acquisition interface, and then click the square "open/open" button under the safety gate on the BSRF safety interlock control system. You can hear the sound of the valve starting, and the warning lights on the FOE and Hutch light up. At this time, the exposure begins. (10) Exposure ends. When the value under "Current and mAH" in the lower right corner of the LIGA light intensity collection interface is greater than or equal to the set exposure, the exposure will automatically stop. The square "open" button light under the safety shutoff on the BSRF safety interlock control system automatically goes out, and the red square "close" button below automatically lights up. The police lights on the FOE and Hutch automatically go off. (11) Close the gate valve. Press the "open" button of the gate valve to close the gate valve. (12) Stop scanning. Press the "stop button" on the touch screen to stop scanning, and then press the "return to origin" button, the scanning device will automatically return to the origin. After reaching the origin, the "Return to Origin" button will turn green and display "Return to Origin". (13) Open the Hutch door. Press the "Open Door" button on the control box outside the door to open the Hutch door. (14) Deflate. First, unscrew the fixing screw of the exposure chamber door and open the fixed door bolt. About half a minute after closing the gate valve, press the "open" button of the vent valve to deflate the exposure chamber. The sound of deflation can be heard during the deflation process. (15) Take samples. After the deflation is completed, open the door of the hard-line exposure chamber, pull out the sample holder, and then close the door of the exposure chamber. Remove the sample from the sample holder. (16) If you want to change the sample and continue the exposure, start the cycle from step (2). If you no longer want to continue to expose, press the "POWER" button of the water cooler to turn off the water cooler, and press the plug-in button on the water cooler to turn off the external circulation pump. Press the "ON" button of the mechanical pump, and then press the "ON" button of the foreline valve to evacuate the exposure chamber. After about 2 minutes, press the "on" button of the foreline valve, and then press the "off" button of the mechanical pump to turn off the mechanical pump.
1B3-X-ray lithography experimental station, soft line exposure operation steps | Soft line exposure operation steps (1) Turn on the cooling water. Turn on "POWER" on the water cooler panel, press the "COOL" button after about a minute, and then press "PUMP". (2) Place the sample. The sample is fixed in the sample holder, insert the sample holder into the scanning frame in the soft line exposure chamber, and tighten the screws to fix it. Close the exposure chamber door. (3) Start scanning. On the touch screen soft line exposure control system, set the "speed setting" and "swing setting" according to your own experimental needs (note that the parameters after "center height" cannot be changed). After the settings are completed, click "Scan Start" to ensure that the scan starts normally before proceeding to the next step. (4) Hutch door interlocking. Press the square green "Search" button inside the Hutch door, and the door warning light will flash and beep. Press the "Close" button outside the Hutch door and the door will automatically close. After the door is closed, click the square green "Search" button outside the door. After the beep ends, the round green light (Enable) under the safety gate on the safety interlock control system panel will light up. (5) Turn on the mechanical pump. On the soft line control power supply box, press the "on" button of the mechanical pump, and the mechanical pump will start. Then press the "on" button of the front-end valve, and you can hear a loud pumping sound, which will gradually become smaller. (6) Turn on the molecular pump. Evacuate to 14Pa, press the "on" button of the molecular pump, start the molecular pump, and press the "on" button of the ionizer switch. (7) Open the pneumatic valve. When the vacuum is higher than 1.0X10-2Pa, press the "Open" button of pneumatic valve 1, then press the "Open" button of pneumatic valve 2, and open the two pneumatic valves in sequence. (8) Set exposure. On the LIGA light intensity acquisition software interface, enter the exposure amount in the box below "Target current value mAH", then click on the right-facing hollow arrow in the upper left corner, which changes to a solid arrow after clicking. (9) Start exposure. Click the "Start" button on the LIGA light intensity acquisition interface, and then click the square "open/open" button under the safety gate on the BSRF safety interlock control system. You can hear the sound of the valve starting, and the warning lights on the FOE and Hutch light up. At this time, the exposure begins. (10) Exposure ends. When the value under "Current and mAH" in the lower right corner of the LIGA light intensity collection interface is greater than or equal to the set exposure, the exposure will automatically stop. The square "open" button light under the safety shutoff on the BSRF safety interlock control system automatically goes out, and the red square "close" button below automatically lights up. The police lights on the FOE and Hutch automatically go off. (11) Close the pneumatic valve. Press the "Open" button of pneumatic valve 2, then press the "Open" button of pneumatic valve 1 to close both pneumatic valves in sequence. (12) Stop scanning. Press the "stop button" on the touch screen to stop scanning, and then press the "return to origin" button, the scanning device will automatically return to the origin. After reaching the origin, the "Return to Origin" button will turn green and display "Return to Origin". (13) Stop the molecular pump and mechanical pump. Press the "on" button of the power switch. Press the "ON" button of the molecular pump, and the molecular pump speed begins to decrease. When the molecular pump speed drops to 0, press the "ON" button of the foreline valve to close the foreline valve, and then press the "OFF" button of the mechanical pump to close the mechanical pump. (14) Open the Hutch door. Press the "Open Door" button on the control box outside the door to open the Hutch door. (15) Deflate. First, unscrew the fixing screw of the exposure chamber door and open the fixed door bolt. Press the "open" button of the deflation valve to deflate the exposure chamber. The sound of deflation can be heard during the deflation process. (16) Take samples. After the deflation is completed, open the door of the cord exposure chamber, pull out the sample holder, and then close the door of the exposure chamber. Remove the sample from the sample holder. (17) If you want to change the sample and continue the exposure, start the cycle from step (2). If you no longer want to continue to expose, press the "POWER" button of the water cooler to turn off the water cooler. Press the "ON" button of the mechanical pump, and then press the "ON" button of the foreline valve to evacuate the exposure chamber. After about 3 minutes, press the "on" button of the foreline valve, and then press the "off" button of the mechanical pump to turn off the mechanical pump.
1B3-X-ray lithography experimental station, line station status | normal operation
1B3-X-ray lithography experimental station, Beijing synchrotron radiation experiment application process | Step 1: Project application, ● Ordinary project application: Users register and log in to the Chinese Academy of Sciences Major Science and Technology Infrastructure Sharing Service Platform (http://lssf.cas.cn/) (hereinafter referred to as the platform) to submit a project application, then download and print the application form, sign it at the designated location and stamp it with the official seal of the unit, and mail one copy to the BSRF User Office (hereinafter referred to as the user office). The User Office accepts project applications throughout the year. ● Application for key topics: Key topics must meet the following conditions: basic research or applied basic research with significant academic value; applied research that is expected to achieve significant economic benefits or social impact in the short term (usually three years); especially research projects that have received major or key national support, or research that is current research hotspots and frontiers; applicants have good research performance in related fields, have experience in synchrotron radiation work, and have a stable research team. Users who have carried out experiments at BSRF for more than two years can apply for installation key projects. Users who want to apply for key topics also need to submit an application online through the platform, and mail a copy of the application signed and stamped with the official seal of the unit to the user office. The specific reception time for key project applications is subject to the notice issued by the user office, which is generally once a year. Step 2: Project review● Preliminary review of the experimental station: evaluate the feasibility of the experimental technology;● Two peer experts review: scientific significance and academic importance (sent twice for review in March and September each year);● After passing the review of general topics, project management (valid for two years); key topics also need to undergo on-site report defense. After being reviewed and approved by the professional committee and user committee, project management (valid for two years) will be given priority. Step 3: Apply for opportunities and arrange opportunities. ● Initial review of the experimental station: assess the feasibility of the experimental technology; ● Review by two peer experts: scientific significance and academic importance (submitted twice for review in March and September each year); ● After passing the review of general topics, project management (valid for two years); key topics require on-site report defense. After review and approval by the professional committee and user committee, project management (valid for two years) will be given priority. Step 4: Complete safety training and conduct experiments● First-time experimenters need to register an account on the platform and complete personal radiation protection safety training. Only after passing the exam can they enter the experimental hall (the safety training is valid for 2 years);● After entering the experimental line station, users must strictly Comply with the various rules and regulations of BSRF; fill in the light consumption record form, conduct online station safety operation training under the guidance of online station personnel, and then carry out experiments; ● After the experiment, users need to clean the experimental bench and dispose of experimental waste in accordance with relevant regulations; return the access card and personal dose card. Step 5: Conclusion and submission of results ● When publishing user scientific and technological achievements, if all or part of the experimental content is completed on BSRF, it must be noted that "data were obtained at Beijing Synchrotron Radiation Facility XX Station", and relevant achievement information (including articles, awards, patents, social benefits, etc.) should be fed back to the user office in a timely manner.
4W1A-X-ray Imaging Experiment Station, Experiment Manual | 4W1A Beamline and X-ray Imaging Experiment Station.pdf
4W1A-X-ray imaging experimental station, introduction | The X-ray imaging experimental station is located at the end of the 4W1A beam line of the Beijing Synchrotron Radiation Facility. It is an experimental station for both white light and monochromatic light. In the more than ten years since the establishment of the 4W1A beamline and experimental station, it has been mainly used to analyze and study microstructural defects on the surface and inside of crystal materials and devices. After 2000, with the development of X-ray imaging technology, the beamline station became the base for domestic hard X-ray phase contrast imaging research, playing an increasingly important role in phase contrast imaging research in the fields of biology, medicine and materials. The experimental station has also built an X-ray nano-resolution three-dimensional imaging microscope. Using standard samples, the nanomicroscope can obtain spatial resolutions up to 30 nanometers (the spatial resolution obtained by actual samples varies with sample characteristics), which can provide advanced research methods for many disciplines such as life sciences, energy sciences, materials science, microelectronics industry, microprocessing technology and nanotechnology.
4W1A-X-ray imaging experimental station, experimental technology | The main experimental methods carried out include crystal white light topography, micrometer resolution imaging and zone plate nanometer resolution full field imaging and corresponding CT experimental methods. In addition, the experimental station can also carry out research on phase contrast imaging methods, including diffraction enhanced imaging and coaxial phase contrast imaging.
4W1A-X-ray imaging experimental station, research area | life sciences
4W1A-X-ray imaging experimental station, research area | energy science
4W1A-X-ray imaging experimental station, research area | Materials Science
4W1A-X-ray imaging experimental station, research area | Microelectronics industry
4W1A-X-ray imaging experimental station, research area | Micromachining technology
4W1A-X-ray imaging experimental station, research area | Nanotechnology
4W1A-X-ray imaging experimental station, research content | Conduct analysis and research on microstructural defects on the surface and inside of crystal materials and devices.
4W1A-X-ray imaging experimental station, experimental technology | X-ray nanoresolution 3D imaging
4W1A-X-ray imaging experimental station, experimental technology | white light topography
4W1A-X-ray imaging experimental station, experimental technology | X-ray micron resolution imaging
4W1A-X-ray imaging experimental station, experimental technology | Phase contrast imaging method
4W1A-X-ray imaging experimental station, experimental technology | Diffraction enhanced imaging
4W1A-X-ray imaging experimental station, experimental technology | On-axis phase contrast imaging
X-ray nano-resolution three-dimensional imaging, research content | Nano-resolution three-dimensional imaging research
Phase contrast imaging method, research content | Research on the principles and methods of phase contrast three-dimensional imaging;
4W1A-X-ray imaging experimental station, equipment parameters | light source The light source of the 4W1A beam line is a single-period wiggler magnet, and the magnetic field strength of the light source magnet is 1.8T. Because the light source magnet is the first wiggler wiggler magnet in the 4th arc area of the electron storage ring, it is referred to as the 4W1 wiggler. When the electron beam passes through the magnetic field, it is subjected to the Lorentz force to produce circular acceleration motion, and synchrotron radiation X-rays are emitted from the tangential direction of the electron orbit. In the dedicated operation mode of synchrotron radiation (2.5GeV@250mA), the light source parameters of 4W1 (calculated based on coupling coefficient 10%) are: Horizontal bundle size: sx=1.23mm Bundle size in vertical direction: sy=0.42mm Horizontal angular divergence: sx'=0.190mrad Vertical angular divergence: sy'=0.131mrad
4W1A-X-ray imaging experimental station, main equipment | Dual crystal monochromator
4W1A-X-ray imaging experimental station, main equipment | Multi-dimensional motion translation and rotation sample stage
4W1A-X-ray imaging experimental station, main equipment | imaging detector
4W1A-X-ray imaging experimental station, main equipment | Nanoresolution imaging equipment
Dual crystal monochromator, introduction | The main function of the dual-crystal monochromator is to monochromatize broad-spectrum white light X-rays into monochromatic X-rays of specific energy. It is the main optical component of the beam line experimental station. The dual-crystal monochromator consists of two perfect silicon single crystals and two sets of high-precision translation rotation stages. The diffraction surface of the silicon single crystal is Si(111), and the crystal plane spacing d111= 3.1353Å. The process of selecting monochromatic X-rays for a monochromator crystal is Bragg diffraction of the crystal diffraction face of X-rays. The relationship between the wavelength of monochromatic X-rays  and the The relationship between ray energy; the selection of monochromator energy can be automatically adjusted using micron resolution imaging control software.
Multi-dimensional motion translation and rotation sample stage, introduction | The motion adjustment system of the experimental station is the key equipment to achieve fine adjustment of the main equipment of the experimental station, including the monochromator, sample stage and detector. It is built with KOHZU high-precision displacement stage and rotation stage. The sample stage adopts a five-dimensional adjustment mechanism, including rotation, horizontal and vertical adjustment of the sample stage, and two-dimensional adjustment of the sample rotation axis. The rotation and translation adjustment accuracy of the sample stage reaches ±0.005 degrees and ±0.3 microns respectively; the detector includes horizontal and vertical two-dimensional adjustments. Using stepper motors and controllers, the experimental station developed a motion adjustment system control program that can realize automatic adjustment of samples and efficient collection of CT data.
Imaging detectors, introduction | Commonly used detectors in experimental stations include X-ray film, ionization chamber and X-ray sCMOS detector. 1) X-ray film The experimental station is equipped with FUJI-IX50 and IX80 industrial X-ray films. The diameter of the film latex particles before development is about 0.3~0.8m. It is used with X-ray development and fixer and is mainly used for white light topography of single crystal materials. 2) Ionization chamber The ionization chamber used in the experimental station is an Ar gas-filled ionization chamber, which is often used to measure the relative light intensity of X-rays when dimming or scanning the crystal rocking curve. • When using, ensure that X-rays are incident on the window of the ionization chamber; • Apply a certain bias voltage (400V) to the two plates of the ionization chamber; • The photocurrent signal measured in the ionization chamber is amplified by a microcurrent meter (KEITHLEY 6517) and transmitted to the computer for data collection and display; 3) X-ray sCMOS detector The commonly used detectors at the experimental station are a set of fiber-coupled sCMOS imaging detectors C12849 from Hamamatsu, Japan, and a set of self-developed lens-coupled imaging detectors with switchable magnifications. The lens-coupled detector can switch 1X, 2X, 4X, and 10X lenses. Hamamatsu Flash4.0 sCMOS and Hamamatsu C15550 qCMOS detectors are commonly used for image signal collection.
X-ray micron resolution imaging, experimental procedures | Through the double crystal monochromator of the experimental station, the white light X-rays at the beam line exit are monochromated, and monochromatic X-rays of a specific energy are selected to be irradiated on the sample. After absorption by the sample, a two-dimensional projection image of the internal microstructure of the sample is obtained on the detector. The optical path diagram of the micron resolution imaging experiment is shown in Figure 2. Due to the advantages of the 4W1 light source such as high brightness, high collimation, small luminous point size, and ultra-long object source distance, high-resolution images can be obtained. For example, a detector using a 10X lens can obtain a spatial resolution of approximately 1m. The experimental station has developed micron-resolution imaging data collection software based on the EPICS control system, which can realize a variety of imaging data collection functions, such as: single image collection, real-time collection; continuous collection with set intervals (suitable for in-situ experiments); CT collection of single energy points and flying CT collection based on external triggering. The software interface is shown in Figure 3.
X-ray nanometer resolution three-dimensional imaging, experimental procedures | The principle of X-ray nano-resolution three-dimensional imaging is shown in Figure 4. Through the zone plate "lens" magnification imaging, the CCD detector cooperates with the precision sample turntable to capture magnified projection images of the sample at different angles, and uses the CT reconstruction algorithm to obtain the three-dimensional density distribution of the sample. This X-ray microscope currently provides two experimental modes: high-resolution mode and large field of view mode: the high-resolution mode has an imaging field of view of 15 μm × 15 μm, and the achievable resolution is 50 nm (standard sample); the large field of view mode has an imaging field of view of 65 μm × 65 μm, and the achievable resolution is 150 nm (standard sample).
White light topography, experimental procedures | white light topography X-ray topography is a method to examine the surface and internal microstructure defects of crystal materials and devices based on the changes in diffraction contrast and image elimination rules in the crystal. It is widely used in the study of the integrity of crystal materials. The white light morphology experiment uses the Laue diffraction method. The incident X-rays are continuous broad-spectrum X-rays. Each Laue spot obtained represents the structural information of the corresponding diffraction surface. The key points in the white light morphology experiment are: • The research object of the white light morphology experiment is single crystal materials; • The white-light X-ray interface size provided by the experimental station is up to 25×14mm2, which can be applied to the research of larger-sized samples; • In white light morphology experiments, a precision four-knife slit can be used to adjust the incident light spot to a suitable size. For some samples with large scattering, the size of the incident light spot should be smaller than the size of the sample. • For direct X-rays transmitted through the sample, lead sheets must be pasted on the corresponding parts of the film box to block them; • White light morphology experiments use X-ray film to capture Laue diffraction spots, and the exposure time of the film is generally a few seconds; • In order to prevent the influence of scattered light on imaging, lead sheets can be used to block stray light in white light morphology experiments; • When preparing samples for white light morphology experiments, the residual stress inside the samples should be reduced as much as possible; • After obtaining the results of the white light morphology experiment, the Laue diffraction spots must be indexed to determine the crystal structure information represented by each diffraction spot.
4W1A-X-ray imaging experimental station, line station status | normal operation
4W1A-X-ray Imaging Experiment Station, Beijing Synchrotron Radiation Experiment Application Process | Step 1: Project application, ● Ordinary project application: Users register and log in to the Chinese Academy of Sciences Major Science and Technology Infrastructure Sharing Service Platform (http://lssf.cas.cn/) (hereinafter referred to as the platform) to submit a project application, then download and print the application form, sign it at the designated location and stamp it with the official seal of the unit, and mail one copy to the BSRF User Office (hereinafter referred to as the user office). The User Office accepts project applications throughout the year. ● Application for key topics: Key topics must meet the following conditions: basic research or applied basic research with significant academic value; applied research that is expected to achieve significant economic benefits or social impact in the short term (usually three years); especially research projects that have received major or key national support, or research that is current research hotspots and frontiers; applicants have good research performance in related fields, have experience in synchrotron radiation work, and have a stable research team. Users who have carried out experiments at BSRF for more than two years can apply for installation key projects. Users who want to apply for key topics also need to submit an application online through the platform, and mail a copy of the application signed and stamped with the official seal of the unit to the user office. The specific reception time for key project applications is subject to the notice issued by the user office, which is generally once a year. Step 2: Project review● Preliminary review of the experimental station: evaluate the feasibility of the experimental technology;● Two peer experts review: scientific significance and academic importance (sent twice for review in March and September each year);● After passing the review of general topics, project management (valid for two years); key topics also need to undergo on-site report defense. After being reviewed and approved by the professional committee and user committee, project management (valid for two years) will be given priority. Step 3: Apply for opportunities and arrange opportunities. ● Initial review of the experimental station: assess the feasibility of the experimental technology; ● Review by two peer experts: scientific significance and academic importance (submitted twice for review in March and September each year); ● After passing the review of general topics, project management (valid for two years); key topics require on-site report defense. After review and approval by the professional committee and user committee, project management (valid for two years) will be given priority. Step 4: Complete safety training and conduct experiments● First-time experimenters need to register an account on the platform and complete personal radiation protection safety training. Only after passing the exam can they enter the experimental hall (the safety training is valid for 2 years);● After entering the experimental line station, users must strictly Comply with the various rules and regulations of BSRF; fill in the light consumption record form, conduct online station safety operation training under the guidance of online station personnel, and then carry out experiments; ● After the experiment, users need to clean the experimental bench and dispose of experimental waste in accordance with relevant regulations; return the access card and personal dose card. Step 5: Conclusion and submission of results ● When publishing user scientific and technological achievements, if all or part of the experimental content is completed on BSRF, it must be noted that "data were obtained at Beijing Synchrotron Radiation Facility XX Station", and relevant achievement information (including articles, awards, patents, social benefits, etc.) should be fed back to the user office in a timely manner.
4W1B-X-ray fluorescence microanalysis experimental station, experimental manual | 4W1B Fluorescent Line Station Software Operation Manual.pdf
4W1B-X-ray fluorescence microanalysis experimental station, research area | Physics
4W1B-X-ray fluorescence microanalysis experimental station, research area | catalysis
4W1B-X-ray fluorescence microanalysis experimental station, research area | biomedicine
4W1B-X-ray fluorescence microanalysis experimental station, research area | environmental science
4W1B-X-ray fluorescence microanalysis experimental station, research area | Materials Science
4W1B-X-ray fluorescence microanalysis experimental station, research area | Humanities and Archeology
4W1B-X-ray fluorescence microanalysis experimental station, research area | Legal appraisal
4W1B-X-ray fluorescence microanalysis experimental station, research content | 1. Micro-area analysis (μXRF and μXANES) methods can be used to conduct non-destructive composition and valence analysis of micro-areas and trace elements of various materials, including geology and minerals, biomedicine, environmental science, materials science, humanities and archeology, and public safety. 2. Using X-ray emission spectroscopy (XES) experimental methods, research covers 3D transition metal elements, involving catalytic materials, materials, minerals, ecology, metalloproteins and other related fields of chemistry, earth science, environmental science and life science.
4W1B-X-ray fluorescence microanalysis experimental station, introduction | The Beijing Synchrotron Radiation Facility X-ray Fluorescence Microanalysis Experimental Station is located at the end of the 4W1B beam line in Hall 12 of the Beijing Synchrotron Radiation Facility (Figure 1). In 2018, the 4W1B beamline was upgraded to an X-ray pink light micro-focusing beamline, with a working distance of 20mm, a spot size of 50μm at the sample, a luminous flux of about 1013phs/s, and a minimum element detection limit of ppb (μg/g) level. X-ray fluorescence spectrum analysis can be carried out, including conventional micro-area XRF, 2D\3D element distribution, X-ray emission spectroscopy (XES) and other experiments. XRF covers all elements after Na; XES mainly covers 3d transition metals.
4W1B-X-ray fluorescence microanalysis experimental station, experimental technology | Micro X-ray Fluorescence Analysis (μXRF)
4W1B-X-ray fluorescence microanalysis experimental station, experimental technology | X-ray emission spectroscopy (XES)
Micro-area X-ray fluorescence analysis (μXRF), research content | Use flying scan mode mapping to obtain 2D/3D element distribution.
Micro-area X-ray fluorescence analysis (μXRF), research content | Trace element composition and distribution of materials
X-ray emission spectroscopy (XES), research content | Do high-pressure, high-temperature, electro/photocatalytic in-situ ctc-XES and vtc-XES,
X-ray emission spectroscopy (XES), research content | 3d transition metal elements
X-ray emission spectroscopy (XES), research content | Use ctc-XES of transition metals, namely Kβ1,3 and its satellite peak Kβ′, to obtain valence and spin state information.
X-ray emission spectroscopy (XES), research content | Use vtc-XES, namely Kβ2,5 and its satellite peak Kβ″, to study the valence state and ligand information of elements.
X-ray fluorescence holography, research content | Obtain 3D structural information of elements.
4W1B-X-ray fluorescence microanalysis experimental station, spot parameters | In terms of spot parameters at the sample, there are two experimental modes: quasi-monochromatic light focusing mode and pink light focusing mode. Among them, the energy range of the quasi-monochromatic light focusing mode is 8-15keV, the energy resolution (ΔE/E) is less than 7×10⁻², the luminous flux is greater than 1×10¹² photons/s at 15keV, and the spot size (H×V) is 50μm; the energy range of the pink light focusing mode is 5-18keV, the energy resolution data is not mentioned, and the luminous flux is greater than 1×10¹³ photons/s, the spot size (H×V) is also 50μm.
4W1B-X-ray fluorescence microanalysis experimental station, beam line structure | The main optical components in the rear beam line of the front-end area are, in order, water-cooled vertical double-blade slit (Slit1), water-cooled diamond filter, water-cooled dual multilayer monochromator (DMM, Dubble Multilayer Monochromator), toroidal focusing mirror (TM, Toroidal Mirror), high-precision adjustable four-blade slit (Slit2), and Kirkpatrick Baez focusing mirror (K-B Mirror) and multi-capillary half-lens (Multi-capillary half-lens). The first three optical elements are installed in the beamline FOE, while the last three are installed in the experimental station.
4W1B-X-ray fluorescence microanalysis experimental station, main equipment | KOUZU three-dimensional electric scanning table
4W1B-X-ray fluorescence microanalysis experimental station, main equipment | Vortex ME4 solid state detector, SOPHIA-XO surface detector and Pilatus 100k surface detector, diode
4W1B-X-ray fluorescence microanalysis experimental station, main equipment | A Leica long working distance microscope, stereo microscope and conventional visible light sampling camera
4W1B-X-ray fluorescence microanalysis experimental station, experimental sample processing | Contact information for purchasing special tape for fluorescence analysis and slicing experimental samples 1. Special tape purchase information American Canadian Instrument Co., Ltd., model: TF-500, XRF tape 2. Experimental sample section information The only requirement for XRF experiments is that the sample surface is flat. Secondly, in order to reduce the matrix effect (Matrix effect), the sample needs to be sliced, generally with a thickness of about 30-50 microns. The carrier of the section must be clean (containing no elements to be measured, or containing elements below the detection limit) film or tape. We recommend the XRF special tape of American Canadian Instrument Co., Ltd.
4W1B-X-ray fluorescence microanalysis experimental station, data processing software | At present, the mapping software of the experimental station uses 4W1B-XRF1.0 developed by the experimental station itself and PyMCA developed by ESRF.
4W1B-X-ray fluorescence microanalysis experimental station, experimental procedures | 1. When looking for dispenser BeamtimeID (Computer 1): Use a browser to open the website http://4w1b.ihep.ac.cn Login account: bsrf4w1b@ihep.ac.cn The password is obtained from line station personnel or colleagues in the computing center. After logging in, select "Line Station" → "4W1B" → "Machine Time Arrangement" and search for the BeamtimeID corresponding to the experiment based on the proposal name, PI name, and experiment time. 2. Conduct the experiment (Computer 1): 2.1 Start the Mamba backend. Right-click on the desktop and select Open Terminal. Enter as shown below (note the underline): 2.2 Start Mamba Data Worker. Right-click on the desktop and select Open Terminal. Enter as shown below (note the capitalization): 2.3 Start the Mamba front end. Double-click the Mamba icon on the desktop, and the Mamba user interface will be displayed after startup. 2.4 Click Auth-Login as shown below, enter the username, password and BeamtimeID. After successfully logging in, you can conduct the experiment. 2.5 View and fill in metadata information. Click Metadata-metadata generator to open the window below and automatically obtain proposal information. The user manually fills in the sample information and clicks submit to submit the filled-in information. 2.6 Select file storage location (1) Select Parameter Setting in the lower left corner, enter the file storage window, click the open button to open the file storage location interface, and fill in the saved file name in the File Name position. If the file name is empty, a file name in the form of "test+timestamp.nxs" will be generated by default. All files are stored in nxs format by default. If you need to store them in h5 format, you must add h5 or hdf5 suffix to the file name (for example: aaa.h5) (2) Click open to select the current path, and click the Add Folder button under open to add a folder to the current directory. 2.7 Select scanning area (1) Place the sample on the sample holder and adjust motor m3 to make the sample clearer under the microscope. (2) Click View—>Monitor in the menu bar to open the Monitor window. (3) The user can click the Move button on the toolbar to make it selected, and select the position to which the light spot is to be moved in the picture currently displayed in the Monitor window. After the motor stops moving (judged by whether the motor position value in the Motor column of the Mamba main interface changes), click the Count button to view the camera image of the current motor position, which is convenient. (4) Step selection. After this step is completed, it is recommended to click the Move button again to turn it into an unselected state to avoid misoperation. (4) Click the Select Range button on the toolbar (be sure to click the Count button to update the image before clicking the Select Range button). A selection ROI will appear. Move the ROI box, select the scan area, and click the Submit ROI button. You can see the parameters of the scan area in Scan Mechanism, complete other parameters, and start scanning. Special attention should be paid to the fact that the motor position cannot be moved after clicking the Select Range button. If there is any movement, please click Cancel in the lower right corner of the Monitor interface. After canceling, repeat this step (first click the Count button, then click the Select Range button). 2.8 Count operation before collection (1) Count can be performed before the experiment starts. As shown in the figure below, count num (the number of counts, required, if not filled in, the default is 1) and count delay (when two or more counts are performed, the time interval between each two counts, optional), and then click count. If you need to change the exposure time of the detector, you can enter Acquire time. (2) Click PyMca Plot to see the count results. If the user needs to use PyMca for online analysis, PyMca's cfg configuration file generation can be completed in this step (see 4.5 for details). To ensure the synchronization of data collection and analysis, please try to select less than 5 elements. 2.9 Online analysis settings Online analysis currently supports two methods, "Add" and "PyMca" as shown in the figure below. (1) If you perform rough online analysis, select the channel summation mode (Add). The number of summing channels needs to be set. The left side is the smaller value, the right side is the larger value (cannot be greater than 4095). The number of channels can be determined by observing the position of the characteristic peak after executing count. (2) If you want to do precise online analysis, choose PyMca. Select PyMca Fitting, select PyMca, select the cfg file analyzed in step 2.8 (2) in Fit Configuration File, click Apply, and you can see the configured elements in the plot2d window. 2.10 Set scanning parameters. Select the experimental mode on the left and fill in the relevant parameters on the right. Currently, it supports multiple acquisition modes such as Fly, Grid scan snake, and Count. If you want to use Fly Scan, you need to select Fly. When filling in the Fly Scan parameters: 1. When filling in the parameters on the right side, first fill in the Period and Duty. Period is the time required for each step, and Period*Duty is the detector exposure time for each step; 2. First select m2 (vertical axis) and fill in the start and end coordinates and the number of scanning points; then select m1 (horizontal axis) and fill in the parameters. 3. Referring to the data recorded in the chapter about determining the scanning area of ​​the mobile motor, the starting point of m2 is b and the end point is d; the starting point of m1 is a and the end point is c. Just fill it in one by one. 4. Number of scanning points = ((end coordinate - start coordinate)/step size) + 1 2.11 Click the red box button to start collection. 2.12 Secondary selection scanning function. After users perform a rough scan of a certain area, if they want to perform a detailed scan of certain details in the area, they can use the secondary selection scan function. The user clicks the Select Range button in any plot2d control toolbar in the lower right corner to generate an ROI box in the plot2d control. The user drags the ROI box, selects the area to be scanned twice, and uses Point Num to set the number of scan points. Then click the Submit button to send the selected area to the scan control Scan Mechanism. Each time the Submit button is clicked, a new scan will be added after the Scan Mechanism. Repeat 2.10 and 2.11. Steps to complete the second scan. Users can select multiple regions by clicking the Select Range button multiple times in a plot2d control. By double-clicking the ROI region, the ROI selection dialog box can be opened. 2.13 User logout. After the experiment, click Auth-Logout as shown below, and click logout to log out of the Mamba system. 3. View and download experimental data (Computer 1 or user's own computer): 3.1 Use a browser to open the website http://4w1b.ihep.ac.cn Login account: bsrf4w1b@ihep.ac.cn The password is obtained from line station personnel or colleagues in the computing center. 3.2 View and download experimental data. For data with a large amount of data, you can download the client and download it at high speed. 3.2.1 Users can query the required data and download it under the online site -> 4W1B -> data menu; this page only displays the original data in the /raw experiment directory. 3.2.2 Users can view and download the required data according to the data storage directory under the online site -> 4W1B -> File Management menu; on this page, users can search and download the temporary data after data processing in the /processed, /scratch and /share experimental directories according to the data storage location.
4W1B-X-ray fluorescence microanalysis experimental station, 4W1B-XRF1.0 software for data processing | Data processing (PC 2 or user's own computer): 1) Use 4W1B-XRF1.0 software for data processing Due to the fast speed of flying scanning, the statistics of fluorescence peaks of some elements are not enough, and the effect of fitting spectra using PyMCA is not good. Therefore, 4W1B cooperated with colleagues in the software group to develop 2D/3D diagrams generated from experimental data, as well as element correlation analysis software packages 4W1B￾XRF1.0A and B. This software is suitable for flying scan mode and situations where the element content is too low and peak statistics are insufficient. Instructions for using experimental data to generate 2D graphs are as follows: A. View and download experimental data: Use a browser to open the website http://4w1b.ihep.ac.cn Login account: bsrf4w1b@ihep.ac.cn Password: Bsrf4w1b@2022 After logging in, select "Line Station" → "4W1B" → "Machine Time Arrangement" and search for the BeamtimeID corresponding to the experiment based on the proposal name, PI name, and experiment time Select "Line Station" → "4W1B" → "Data", enter the BeamtimeID corresponding to the experiment to view and download the experimental data. For data with a large amount of data, you can download the client and perform high-speed downloading
4W1B-X-ray fluorescence microanalysis experimental station, 4W1B-XRF1.0A for data processing | Using 4W1B-XRF1.0A for data processing: This program combines the channel addition program and the oversampling deconvolution program, including an executable program (ChannelAddPlusDeconv_Ar_Normalization.exe) and configuration file (config.ini). Users need to put these two files under the same folder. After modifying the corresponding parameters in the configuration file, double-click the executable file to process the data and obtain the final element distribution map and output data file. Modify the corresponding parameters in the configuration file There are corresponding parameter descriptions in the configuration file: After setting the parameters, save and close the configuration file. Double-click the executable program, wait for a moment, and the picture will pop up (Figure 1, Figure 2 and Figure 3) These three pictures are the results of data processing after channel summation: Figure 1: With axes and color bar Figure 2: Only color bar, no axis Figure 3: None of the above, and this image is used as the input image for deconvolution After closing Figure 3, the picture Figure 1 pops up again (the left picture is the deconvolution input picture, the right picture is the deconvolution output picture) After closing the image, all txt files and png files will be generated in this folder
4W1B-X-ray fluorescence microanalysis experimental station, PyMca software for data processing | Data processing using PyMca software For data where the new signal-to-noise ratio of the spectral peak is good enough and the user wishes to do quantitative analysis, PyMCA can be used to process it. Specific data processing can be done through computers on the online site, using the computing resources provided by the Institute of High Energy Computing Center for online data processing (faster speed). Users can also install PyMca software on their own computers for offline data processing. 4.0 If you need to use the computing resources provided by the Institute of High Energy Computing Center for online data processing, you first need to log in to the computing center server on computer 2, open the terminal and log in to the 192.168.68.16 server through user bsrf4w1b, that is, enter the following command: $ ssh bsrf4w1b@192.168.68.16 -X Enter the password (please contact colleagues in the computing center to obtain the password) to log in successfully. At this time, the beginning of the line of the input command is displayed. (base) word, enter $ pymca Then start the PyMca software. If you process data on your own computer, you can download and install the PyMca software yourself: https://sourceforge.net/projects/pymca/ 4.1 Import raw data (.h5 or .nxs format) When the computing resources provided by the Computing Center of the Institute of High Energy are used for online data processing, experimental data need to be found from the data storage server. User data is stored in /hepsfs/bl/4W1B/202210/Data/beam-time-id/raw/ (hereinafter referred to as the raw folder) For example: /hepsfs/bl/4W1B/202210/Data/GB06-20221010-01/raw/ Note: The bold red part is an example of BeamtimeID. Different users have their own BeamtimeID, and they need to confirm that the BeamtimeID in the path is their corresponding ID. The input file type can be selected via Files of type, which can be set to All Files(*). 4.2 Expand the data and double-click xpress_sum_merge 4.3 Check Signals, select Force MCA, click ADD, and the one-dimensional spectrum will appear on the right 4.4 Energy calibration 4.5 Fit and generate .cfg files (only some basic configurations are shown here, see the fifth edition of the Fluorescence Station Data Processing Tutorial for details). It should be noted that when using the computing resources provided by the Institute of High Energy Computing Center for online data processing, not all locations can store .cfg files. Only /hepsfs/bl/4W1B/202210/Data/beamtime-id/processed/ (hereafter referred to as the processed folder) has write permissions and the .cfg file must be stored in this path. 4.6 Batch processing (1) Select Tools→Batch Fitting in the upper menu bar to enter the batch processing interface (here, we use the Institute of High Energy Computing Center for online data processing as an example. The native PyMca software does not have the image display function on the right). (2) Click Browse at the top, select the input file in the raw folder, and select the input file type through Files of type, which can be set to All Files (*). Expand the data and double-click on xpress_sum_div_I0. Data-related information will appear below. Check the box as shown below and click Done. (3) Then select the .cfg file path and output file path (both in the processed folder) (4) When using the Institute of High Energy Computing Center for online data processing, you can increase the processing speed by checking Use several processes. You need to set the value of Number of processes (within 100). Note: The larger the value of processes, the faster the processing speed. However, too many processes will affect the time of the progress bar pop-up window and the time of displaying the final drawing data. Just wait a few seconds. (5) Click start to start batch processing, and a progress bar window will pop up. (You may need to wait a few seconds, do not click the start button repeatedly) (6) After running, a series of files will be generated in the output file path, among which .DAT files can be used for drawing. For drawing methods using Origin, please refer to the fifth edition of the Fluorescence Station Data Processing Tutorial. It should be noted that the motor position is stored in The last two columns in /entry/instrument/pandabox in the original data. When viewing the data, right-click the pandabox, select show information, and then view the data in the DataView column of the right window in the figure below. Note the settings in the green box at the bottom of the image. In addition, the xpress_sum_div_I0 data used in batch processing has been normalized using ionization chamber counts. There is no need for users to repeat normalization and batch processing can be performed directly. The 4th column data in the blue box in the figure below is the ionization chamber count for user reference. (7) When using the Institute of High Energy Computing Center for online data processing, the table on the right will automatically load element data after the operation is completed, and the corresponding elements can be checked to draw images. a) Check select all to draw an image in which all elements overlap. Click again to uncheck and the image will be cleared. Note: The calculation of select all will be slower, and you need to wait a few seconds for the graph to appear. b) Check one or more elements to draw the image. c) You can follow the mouse below to display the horizontal and vertical axes and corresponding pixel value data. d) The current image can be saved (can only exist in the processed folder)
4W1B-X-ray fluorescence microanalysis experimental station, common faults | 1. The detector IOC fails, the Mamba backend reports an error as shown below, and the Mamba frontend is stuck: Follow these steps to restore the software: (1) Close the Mamba front-end, terminate the program with ctrl+c in the terminal window of the Mamba Data Worker, enter exit in the python command line in the terminal window of the Mamba back-end, and press Enter to exit the Mamba back-end; (2) Open a new terminal and run the following command to restart the detector IOC: $ ioc-stop xsp3 $ ioc-start xsp3 (3) At this time, follow Part 3 of this document to restart Mamba and continue the experiment
4W1B-X-ray fluorescence microanalysis experimental station, line station status | normal operation
4W1B-X-ray fluorescence microanalysis experimental station, Beijing synchrotron radiation experiment application process | Step 1: Project application, ● Ordinary project application: Users register and log in to the Chinese Academy of Sciences Major Science and Technology Infrastructure Sharing Service Platform (http://lssf.cas.cn/) (hereinafter referred to as the platform) to submit a project application, then download and print the application form, sign it at the designated location and stamp it with the official seal of the unit, and mail one copy to the BSRF User Office (hereinafter referred to as the user office). The User Office accepts project applications throughout the year. ● Application for key topics: Key topics must meet the following conditions: basic research or applied basic research with significant academic value; applied research that is expected to achieve significant economic benefits or social impact in the short term (usually three years); especially research projects that have received major or key national support, or research that is current research hotspots and frontiers; applicants have good research performance in related fields, have experience in synchrotron radiation work, and have a stable research team. Users who have carried out experiments at BSRF for more than two years can apply for installation key projects. Users who want to apply for key topics also need to submit an application online through the platform, and mail a copy of the application signed and stamped with the official seal of the unit to the user office. The specific reception time for key project applications is subject to the notice issued by the user office, which is generally once a year. Step 2: Project review● Preliminary review of the experimental station: evaluate the feasibility of the experimental technology;● Two peer experts review: scientific significance and academic importance (sent twice for review in March and September each year);● After passing the review of general topics, project management (valid for two years); key topics also need to undergo on-site report defense. After being reviewed and approved by the professional committee and user committee, project management (valid for two years) will be given priority. Step 3: Apply for opportunities and arrange opportunities. ● Initial review of the experimental station: assess the feasibility of the experimental technology; ● Review by two peer experts: scientific significance and academic importance (submitted twice for review in March and September each year); ● After passing the review of general topics, project management (valid for two years); key topics require on-site report defense. After review and approval by the professional committee and user committee, project management (valid for two years) will be given priority. Step 4: Complete safety training and conduct experiments● First-time experimenters need to register an account on the platform and complete personal radiation protection safety training. Only after passing the exam can they enter the experimental hall (the safety training is valid for 2 years);● After entering the experimental line station, users must strictly Comply with the various rules and regulations of BSRF; fill in the light consumption record form, conduct online station safety operation training under the guidance of online station personnel, and then carry out experiments; ● After the experiment, users need to clean the experimental bench and dispose of experimental waste in accordance with relevant regulations; return the access card and personal dose card. Step 5: Conclusion and submission of results ● When publishing user scientific and technological achievements, if all or part of the experimental content is completed on BSRF, it must be noted that "data were obtained at Beijing Synchrotron Radiation Facility XX Station", and relevant achievement information (including articles, awards, patents, social benefits, etc.) should be fed back to the user office in a timely manner.
4B7A-Zhongneng Experiment Station, Experiment Manual | 4B7A 4B7B beamline general operations.pdf
4B7A-Zhongneng Experiment Station, Introduction | Line station 4B7A is led out by bent iron and shares a front-end area with soft X-ray beam line station 4B7B. This line station provides medium-energy X-rays (2.1~6keV), covering the K absorption edge of low-Z elements (P, S, Cl, K, Ca, Ti, etc.) and the L absorption edge of some elements. It can provide absorption spectrum measurement services in partial fluorescence yield (PFY) mode and electron yield (TEY) mode for users in scientific fields such as energy environment, materials, chemistry, biology, geophysics, etc. In addition, the line station also provides the necessary infrastructure to meet users' in-situ experimental needs.
4B7A-Zhongneng Experimental Station, research area | Energy environment
4B7A-Zhongneng Experimental Station, research area | Material
4B7A-Zhongneng Experimental Station, research area | Chemistry
4B7A-Zhongneng Experimental Station, research area | biological sciences
4B7A-Zhongneng Experimental Station, research area | Detector calibration
4B7A-Zhongneng Experimental Station, research area | Optical component characterization
4B7A-Zhongneng Experimental Station, research area | X-ray radiation measurement standards
4B7A-Zhongneng Experiment Station, Experimental Technology | X-ray absorption spectroscopy (XAS)
4B7A-Zhongneng Experiment Station, Experimental Technology | X-ray detector standard transfer and calibration
4B7A-Zhongneng Experiment Station, Experimental Technology | Transmittance of optical components
4B7A-Zhongneng Experiment Station, Experimental Technology | X-ray optical metrology standard determination
4B7A-Zhongneng Experiment Station, Experimental Technology | Reflectance calibration
4B7A-Zhongneng Experiment Station, Experimental Technology | Diffraction efficiency calibration
4B7A-Zhongneng Experiment Station, Experimental Technology | Near-edge X-ray absorption spectroscopy (XANES)
X-ray absorption spectroscopy (XAS), research content | Sample structure and composition
X-ray detector standard transfer and calibration, research content | X-ray detector standard transfer and calibration
Transmittance of optical components, research content | Transmittance of optical components
X-ray optical metrology standard determination, research content | X-ray optical metrology standard determination
Reflectance calibration, research content | Reflectance calibration
Diffraction efficiency calibration, research content | Diffraction efficiency calibration
Near-edge X-ray absorption spectroscopy (XANES), research content | Analyze the valence state of elements
Near-edge X-ray absorption spectroscopy (XANES), research content | Compound composition and empty state electronic structure near the Fermi level.
4B7A-Zhongneng Experiment Station, light spot parameters at the sample | Spot parameters at sample Energy range: Si(111): 2100eV-5700eV, InSb(111): 1750eV-3400eV Energy resolution (E/ΔE): ~7000 Luminous flux (photons/s): ~1x1011 @ S K edge Spot size (HxV): 5x3mm2
4B7A-Zhongneng Experiment Station, light on and off operation | On and off light operation Turning on the light refers to letting synchrotron radiation X-rays irradiate the sample. The steps to open the beam line are as follows: 1. Open the "No. 4 gate valve" (press the dark green button, the red indicator light of the valve goes out and the green indicator light turns on) 2. Open the "safety shutter" (press the green button, the red button light goes out, and the green button light turns on). Turning off the light means to stop synchrotron radiation X-rays from irradiating the sample or entering the sample chamber. The steps to turn off the beam line are as follows: 1. Close the "safety shutter" (press the red button, the green button light goes out, and the red button light turns on) 2. Close the "No. 4 gate valve" (press the dark red button, the green indicator light of the valve goes out and the red indicator light turns on) Precautions for turning on and off the light: 1. The safety shutter should not be opened and closed frequently. It should be closed when the user does not use the beam line for a long time (such as eating, storage ring failure, etc. for more than half an hour). 2. Operate accurately and do not touch other switches
4B7A-Zhongneng Experiment Station, He measurement room electron yield absorption spectrum measurement experimental steps | He Measurement Room Electronic Yield Absorption Spectrum Measurement Operating Procedures 1. Press the 6517B galvanometer "LOCAL" and "Z-CHK" keys to display "ZeroCheck" 2. Deflate: a) Close the No. 4 gate valve and close the safety shutter if you leave for a long time b) Close the angle valve of the He gas measuring chamber c) Open the nitrogen valve switch (turn 90 degrees counterclockwise) d) Open the CF16 fine-tuning valve to deflate e) Observe the air pressure gauge on the top of the He gas measuring chamber. The pointer points to a position close to 0. Close the nitrogen valve switch and the CF16 fine-tuning valve. 3. Change samples: a) Open the hinged door with window b) Remove the alligator clips, loosen the fastening screws and take out the sample holder. Note: Please be sure to wear gloves when operating; be careful of touching the fluorescent probe! ! ! c) Place the sample holder, tighten the fastening screw, and clamp the alligator clip on the sample holder. d) Close the hinged door with window and tighten the nut. 4. Vacuum: a) Make sure to close the CF16 fine-tuning valve b) Turn off the power supply of the molecular pump valve and turn on the power supply of the side pump valve. c) Slowly open the angle valve of the He gas measuring chamber. When the display in the upper right corner of the cabinet begins to decrease, quickly open the angle valve fully. d) When the number displayed in the upper right corner of the cabinet is less than 10, turn off the power supply to the side pump valve and turn on the power supply to the molecular pump valve. e) The experiment can be carried out after the molecular pump works normally for 10 minutes. 5. Experiment: a) Press the "LOCAL" and "Z-CHK" keys of the 6517B ammeter in sequence to display the current value. Note: If the displayed value is unstable or displays "overflow", press the "Z-CHK" key again to display "ZeroCheck", and then contact the person on duty to solve the problem. b) Observe the cabinet vacuum gauge. If it is normal, open the No. 4 gate valve. c) Open the safety shutter d) "He three-phase motor fixed length_sub.vi" adjusts and switches the sample position. e) Collect an absorption spectrum using the TEY program. Note: 1. When changing samples or not conducting experiments for a short period of time, please close the No. 4 gate valve. It is not necessary to close the safety shutter. 2. If you are away for a long time or not conducting experiments, please close the No. 4 gate valve and safety shutter. 3. It is strictly prohibited to leave the experimental station unattended for a long time during the consecration period! 4． Please wear gloves (plastic, rubber or nitrile gloves) when handling samples and sample holders. Bare hands are strictly prohibited! ! ! 5. Do not have special characters in the file name, such as "?:*/\><¦" 6. If you have any questions, please notify the person on duty promptly.
4B7A-Zhongneng Experimental Station, He measurement room fluorescence absorption spectrum measurement experimental steps | He Measurement Room Fluorescence Absorption Spectrum Measurement Operating Procedures 6. Deflate: a) Close the No. 4 gate valve and close the safety shutter if you leave for a long time b) Close the angle valve of the He gas measuring chamber c) Open the nitrogen valve switch (turn 90 degrees counterclockwise) d) Open the CF16 fine-tuning valve to deflate e) Observe the air pressure gauge on the top of the He gas measuring chamber. The pointer points to a position close to 0. Close the nitrogen valve switch and the CF16 fine-tuning valve. 7. Change samples: a) Open the hinged door with window b) Loosen the fastening screws and take out the sample holder. Note: Please be sure to wear gloves when operating; be careful of touching the fluorescent probe! ! ! c) Place the sample holder and tighten the fastening screws. Note: The surface of the sample holder is at a 45-degree angle to the incident light d) Close the hinged door with window and tighten the nut. 8. Vacuum: a) Make sure to close the CF16 fine-tuning valve b) Turn off the power supply of the molecular pump valve and turn on the power supply of the side pump valve; c) Slowly open the angle valve of the He gas measuring chamber. When the display in the upper right corner of the cabinet begins to decrease, quickly open the angle valve fully. d) When the number displayed in the upper right corner of the cabinet is less than 10, turn off the power supply to the side pump valve and turn on the power supply to the molecular pump valve. e) There are no problems with the vacuum and equipment, and experiments can be carried out 9. Experiment: a) Turning on the power of the fluorescence detector must be completed by the person on duty b) Observe the cabinet vacuum gauge. If it is normal, open the No. 4 gate valve. c) Open the safety shutter d) "He three-phase motor fixed length_sub.vi" adjusts and switches the sample position. e) "Energy Scan_SDD_SCA_FY*.vi" collects absorption spectra.
4B7A-Zhongneng Experimental Station, Experimental Procedures for Fluorescence Absorption Spectrum Measurement of Liquid Samples | Liquid sample fluorescence absorption spectrum measurement operating procedures 1. Preparation (this operation is performed by the personnel on duty at the experimental station!): a) Turn off the No. 4 gate valve, the angle valve of the He gas measuring chamber and the power supply of the molecular pump unit in sequence (press the "Work" button, the indicator light next to it goes out) b) Open the nitrogen valve switch (turn 90 degrees counterclockwise) c) Open the CF16 fine-tuning valve to release nitrogen d) Observe the barometer on the top of the He gas measuring chamber until the pointer points to the 0 position above, and close the nitrogen valve switch. e) After the molecular pump stops, turn off the power of the mechanical pump. Remove the quick release connecting the molecular pump (disconnect the compression bellows from the molecular pump) f) Connect the helium valve to the experimental chamber g) Confirm that the fluorescence detector is powered on and warmed up for 15 minutes. 2. Change sample: a) Open the hinged door with window b) Loosen the fastening screws and take out the sample holder. Note: Please be sure to wear gloves when operating; be careful of touching the fluorescent probe! ! ! c) Place the sample holder and tighten the fastening screws. Note: The surface of the sample holder is at an angle of 45 degrees to the incident light. d) Close the hinged door with window and tighten the nut. 3. Helium filled: a) Make sure to open the CF16 fine-tuning valve, and then open the angle valve of the He gas measuring chamber b) Turn on the helium valve switch. c) Open the No. 4 gate valve and adjust the position of the sample so that the light spot hits the sample d) Move the monochromatic light energy to 3.25keV, then open the program Prospect, click "start run", and observe the Ar fluorescence peak of the fluorescence spectrum at 3.0keV e) Click "stop run" and then "start run", and repeat the operation until the Ar fluorescence peak becomes very small, indicating that the air has been eliminated. Click "stop run" to close Prospect f) Close the helium valve and the CF16 fine-tuning valve in sequence, and keep the angle valve of the He gas measurement chamber open; 4． Experiment: f) "He three-phase motor fixed length_sub.vi" adjusts and switches the sample position. g) Collect absorption spectra using the fluorescence program.
4B7A-Zhongneng Experiment Station, precautions for copying experimental data | Notes on copying experimental data: 1. It is strictly prohibited to connect any data storage media (U disk, mobile hard disk, mobile phone, etc.) to the industrial computer that controls the beam line and data collection. 2. After copying out their own data, users have the right to delete their own experimental data on the industrial computer. 3. The operation and maintenance personnel of the experimental station are not responsible for any adverse consequences caused by users not deleting their own data.
4B7A-Zhongneng Experimental Station, Beijing Synchrotron Radiation Experiment Application Process | Step 1: Project application, ● Ordinary project application: Users register and log in to the Chinese Academy of Sciences Major Science and Technology Infrastructure Sharing Service Platform (http://lssf.cas.cn/) (hereinafter referred to as the platform) to submit a project application, then download and print the application form, sign it at the designated location and stamp it with the official seal of the unit, and mail one copy to the BSRF User Office (hereinafter referred to as the user office). The User Office accepts project applications throughout the year. ● Application for key topics: Key topics must meet the following conditions: basic research or applied basic research with significant academic value; applied research that is expected to achieve significant economic benefits or social impact in the short term (usually three years); especially research projects that have received major or key national support, or research that is current research hotspots and frontiers; applicants have good research performance in related fields, have experience in synchrotron radiation work, and have a stable research team. Users who have carried out experiments at BSRF for more than two years can apply for installation key projects. Users who want to apply for key topics also need to submit an application online through the platform, and mail a copy of the application signed and stamped with the official seal of the unit to the user office. The specific reception time for key project applications is subject to the notice issued by the user office, which is generally once a year. Step 2: Project review● Preliminary review of the experimental station: evaluate the feasibility of the experimental technology;● Two peer experts review: scientific significance and academic importance (sent twice for review in March and September each year);● After passing the review of general topics, project management (valid for two years); key topics also need to undergo on-site report defense. After being reviewed and approved by the professional committee and user committee, project management (valid for two years) will be given priority. Step 3: Apply for opportunities and arrange opportunities. ● Initial review of the experimental station: assess the feasibility of the experimental technology; ● Review by two peer experts: scientific significance and academic importance (submitted twice for review in March and September each year); ● After passing the review of general topics, project management (valid for two years); key topics require on-site report defense. After review and approval by the professional committee and user committee, project management (valid for two years) will be given priority. Step 4: Complete safety training and conduct experiments● First-time experimenters need to register an account on the platform and complete personal radiation protection safety training. Only after passing the exam can they enter the experimental hall (the safety training is valid for 2 years);● After entering the experimental line station, users must strictly Comply with the various rules and regulations of BSRF; fill in the light consumption record form, conduct online station safety operation training under the guidance of online station personnel, and then carry out experiments; ● After the experiment, users need to clean the experimental bench and dispose of experimental waste in accordance with relevant regulations; return the access card and personal dose card. Step 5: Conclusion and submission of results ● When publishing user scientific and technological achievements, if all or part of the experimental content is completed on BSRF, it must be noted that "data were obtained at Beijing Synchrotron Radiation Facility XX Station", and relevant achievement information (including articles, awards, patents, social benefits, etc.) should be fed back to the user office in a timely manner.
4B7A-Zhongneng Experimental Station, line station status | normal operation
4B7B-Soft X-ray Experiment Station, Experiment Manual | 4B7A 4B7B beamline general operations.pdf
4B7B-Soft X-ray Experiment Station, Introduction | The Beijing Synchrotron Radiation Facility (BSRF) 4B7B line station is a soft X-ray beamline covering the energy range of 50~1500eV led by a bent iron light source. It shares a front-end area with the medium-energy beamline 4B7A. The line station can provide users with two types of services: 1. Near-edge absorption spectrum measurement in electron yield (TEY) mode; 3. Absorption spectrum measurement in partial fluorescence yield (PFY) mode; 2. Detector calibration and optical element characterization (reflectance of various mirrors, transmittance of filters, crystal diffraction efficiency, etc.).
4B7B-Soft X-ray Experiment Station, Research Area | Energy environment
4B7B-Soft X-ray Experiment Station, Research Area | Material
4B7B-Soft X-ray Experiment Station, Research Area | Chemistry
4B7B-Soft X-ray Experiment Station, Research Area | biology
4B7B-Soft X-ray Experiment Station, Research Area | X-ray metrology
4B7B-Soft X-ray experimental station, research content | 1. Near-edge absorption spectrum measurement in electron yield (TEY) mode; 3. Absorption spectrum measurement in partial fluorescence yield (PFY) mode; 2. Detector calibration, optical element characterization (reflectance of various mirrors, transmittance of filters, crystal diffraction efficiency, etc.)
4B7B-Soft X-ray Experiment Station, Experimental Technology | Near-edge X-ray absorption spectroscopy (XANES)
4B7B-Soft X-ray Experiment Station, Experimental Technology | X-ray optical metrology standard determination
4B7B-Soft X-ray Experiment Station, Experimental Technology | X-ray detector standard transfer and calibration
4B7B-Soft X-ray Experiment Station, Experimental Technology | Transmittance of optical components
4B7B-Soft X-ray Experiment Station, Experimental Technology | Reflectivity Calibration of Optical Components
Near-edge X-ray absorption spectroscopy (XANES), research content | Analyze the valence state of elements
Near-edge X-ray absorption spectroscopy (XANES), research content | Compound composition and empty state electronic structure near the Fermi level.
X-ray optical metrology standard determination, research content | X-ray optical metrology standard determination
X-ray detector standard transfer and calibration, research content | Transmittance of optical components
Transmittance of optical components, research content | Reflectivity Calibration of Optical Components
Reflectance calibration of optical components, research content | Reflectivity Calibration of Optical Components
4B7B-Soft X-ray experimental station, spot parameters at the sample | Spot parameters at the sample Energy range: 50eV-1700eV Luminous flux (photons/s): ~1x1010 @Fe L edge Energy resolution (E/ΔE) R ~1000 Spot size (HxV): 1mmx0.5mm (focus), up to 10mmx5mm (defocus);
4B7B-Soft X-ray experimental station, light on and off operation | 4B7B Beamline General Operations (required reading for users) On and off light operation Turning on the light refers to letting synchrotron radiation X-rays enter the measurement room. The steps to open the beam line are as follows: 1. Turn on all the gate valve control switches except the leftmost switch on the gate valve control box (a green indicator light on the switch is on); 2. Open the "safety shutter" (press the green button, the red button light goes out, and the green button light turns on); 3. Open the "photon shutter" (press the green button, the red button light goes out, and the green button light turns on); Turning off the light means stopping synchrotron radiation X-rays from entering the measurement room. The steps to turn off the beam line are as follows: 1. Close the "safety shutter" (press the red button, the green button light goes out, and the red button light turns on) 2. Close the "photon shutter" (press the red button, the green button light goes out, and the red button light turns on) 3. Close the second and third switches from the left on the gate valve control box; Precautions for turning on and off the light: 1. The "safety shutter" should not be opened and closed frequently. It should be closed when the user does not use the beam line for a long time (such as eating, storage ring failure, etc. for more than half an hour). 2. Operate accurately and do not touch other switches
4B7B-Soft X-ray Experiment Station, precautions for copying experimental data | Notes on copying experimental data: 4. Without permission, it is strictly prohibited to connect any data storage media (U disk, mobile hard disk, mobile phone, etc.) to the industrial computer that controls the beam line and data collection. 5. Copy the data and contact the person on duty at the experimental station.
4B7B-Soft X-ray experimental station, storage ring status | Storage ring status: 5. There is a storage ring operating status display on the cabinet next to the door of the diffraction station shed, or enter the address http://202.122.38.185:8081/ to view it, or install the mobile APP: 6. When the storage ring is injected and the alarm sound and broadcast are heard, the user should leave the experimental station and experimental hall as soon as possible. 7. When the injection of the storage ring is completed, a broadcast is heard and the user can return to the experimental station to continue the experiment. 8. When it is found that the data collected by the program is 0 or NONE or is abnormal, first check the storage ring status to see if the electron beam current (blue data in the upper left corner) is zero. If it is zero, the user should leave the hall and wait for injection. If it is not zero, check whether the "safety light gate" and "photon light gate" or the gate valve are open. After opening, if there is still no signal, contact the person on duty at the experimental station.
4B7B-Soft X-ray experimental station, operations after the experiment | 7. Perform the "light off" operation; 8. For all 6517B electrometers, press the "LOCAL" and "Z-CHK" keys; 9. Fill in the "Experimental Light Record Form" truthfully, and write the dose card number after the name of the experimenter; 10. Notify the personnel on duty at the experimental station and return the access card and dose card; 11. Take away the garbage generated by conducting experiments;
4B7B-Soft X-ray experimental station, absorption spectrum experimental steps | 4B7B Absorption spectroscopy operating procedures Deflate: 1. Perform the "light off" operation; 2. Turn off the 6517B reading. Press the "LOCAL" and "Z-CHK" keys of the 6517B galvanometer to display "ZeroCheck" 3. Turn off the power of the molecular pump and wait for its speed to return to zero (it may take a long time); 4. Turn off the power supply of the dry pump solenoid valve (you can also turn off the power supply of the dry pump); 5. Loosen the knob of the injection window of the sample chamber and ensure that the pressure in the sample chamber exceeds one atmosphere before the injection window can be opened; 6. Open the nitrogen valve, open the bleed valve on the dry pump, and fill the sample chamber with nitrogen; 7. After the injection window leaks, close the nitrogen valve and the bleed valve on the dry pump; note that during a period of time around July when there is high temperature and humidity, it is recommended to vent nitrogen to prevent moisture from entering the sample chamber, which can significantly reduce the vacuuming time; Replace the device to be calibrated: 1. Open the hinged door with a window (that is, the injection window of the sample chamber); 2. Replace the sample. Note: Please be sure to wear gloves (disposable or nitrile) for operation; it is strictly forbidden to touch components and samples in the vacuum with bare hands; 3. Close the nitrogen valve and the bleed valve on the dry pump; 4. Close the trapdoor with window. Vacuuming: 1. Confirm that the bleed valve is closed; 2. Turn on the power of the dry pump and solenoid valve, and start rough vacuuming of the sample chamber; 3. After 2 minutes when the sound of the dry pump decreases significantly, turn on the power of the molecular pump; 4. After the molecular pump speed reaches the rated speed, turn on the vacuum gauge. The indication should be better than a certain empirical value such as 3E-4Pa, so as to make a preliminary judgment that there is no air leakage; 5. When the vacuum degree of the sample chamber reaches a good level of 1e-6Pa, spectroscopy experiments can be carried out; 6. After performing the "turn on" operation, start the experiment. Question: 1. About the electrometer 6517B: If the 6517B displays "overflow", please press the "Z-CHK" key to make the display "ZeroCheck", check whether the signal line is short-circuited. If the above problem still cannot be solved, contact the person on duty. 2. The vacuum degree of the sample chamber must reach a certain threshold before the beam line can be connected to carry out experiments; the vacuum degree is affected by factors such as the cleanliness of the vacuum components, outgassing of materials or samples, length of exposure to the atmosphere, weather and pumping speed; if the vacuum degree is abnormal, please contact the person on duty at the experimental station. 3. Please ensure that the sample placed in the sample chamber or the external equipment connected to the chamber is clean to prevent contamination of the sample chamber and the upstream beam line! ! !
4B7B-Soft X-ray experimental station, line station status | normal operation
4B7B-Soft X-ray Experiment Station, Beijing Synchrotron Radiation Experiment Application Process | Step 1: Project application, ● Ordinary project application: Users register and log in to the Chinese Academy of Sciences Major Science and Technology Infrastructure Sharing Service Platform (http://lssf.cas.cn/) (hereinafter referred to as the platform) to submit a project application, then download and print the application form, sign it at the designated location and stamp it with the official seal of the unit, and mail one copy to the BSRF User Office (hereinafter referred to as the user office). The User Office accepts project applications throughout the year. ● Application for key topics: Key topics must meet the following conditions: basic research or applied basic research with significant academic value; applied research that is expected to achieve significant economic benefits or social impact in the short term (usually three years); especially research projects that have received major or key national support, or research that is current research hotspots and frontiers; applicants have good research performance in related fields, have experience in synchrotron radiation work, and have a stable research team. Users who have carried out experiments at BSRF for more than two years can apply for installation key projects. Users who want to apply for key topics also need to submit an application online through the platform, and mail a copy of the application signed and stamped with the official seal of the unit to the user office. The specific reception time for key project applications is subject to the notice issued by the user office, which is generally once a year. Step 2: Project review● Preliminary review of the experimental station: evaluate the feasibility of the experimental technology;● Two peer experts review: scientific significance and academic importance (sent twice for review in March and September each year);● After passing the review of general topics, project management (valid for two years); key topics also need to undergo on-site report defense. After being reviewed and approved by the professional committee and user committee, project management (valid for two years) will be given priority. Step 3: Apply for opportunities and arrange opportunities. ● Initial review of the experimental station: assess the feasibility of the experimental technology; ● Review by two peer experts: scientific significance and academic importance (submitted twice for review in March and September each year); ● After passing the review of general topics, project management (valid for two years); key topics require on-site report defense. After review and approval by the professional committee and user committee, project management (valid for two years) will be given priority. Step 4: Complete safety training and conduct experiments● First-time experimenters need to register an account on the platform and complete personal radiation protection safety training. Only after passing the exam can they enter the experimental hall (the safety training is valid for 2 years);● After entering the experimental line station, users must strictly Comply with the various rules and regulations of BSRF; fill in the light consumption record form, conduct online station safety operation training under the guidance of online station personnel, and then carry out experiments; ● After the experiment, users need to clean the experimental bench and dispose of experimental waste in accordance with relevant regulations; return the access card and personal dose card. Step 5: Conclusion and submission of results ● When publishing user scientific and technological achievements, if all or part of the experimental content is completed on BSRF, it must be noted that "data were obtained at Beijing Synchrotron Radiation Facility XX Station", and relevant achievement information (including articles, awards, patents, social benefits, etc.) should be fed back to the user office in a timely manner.
BSRF Beijing synchrotron radiation user topics, experiments, and visit process instructions | # BSRF user topics, experiments, and visit process instructions 1. Submit a topic Method 1: Log in to the Chinese Academy of Sciences' major science and technology infrastructure shared service platform (https://lssf.cas.cn/) (hereinafter referred to as the Academy Platform), select "Beijing Synchrotron Radiation Facility User Project Management" from the common links in the upper right corner of the homepage, or select the Beijing Electron Positron Collider BEPC after adding a new application topic, fill in and submit the project application form according to the process; Method 2: Log in to the BSRF user service system (https://user.bsrf.ihep.ac.cn/), click "Project Application" in the left navigation bar, fill in and submit the project application form according to the application steps. Note: The login account of the BSRF user service system is the email address registered when registering on the hospital platform (please note that it is not the login account of the hospital platform), and the password is the same as the login password of the hospital platform. 2. Reservation time After the subject application is approved, click "Schedule Reservation" in the BSRF user service system, follow the steps to fill in and submit the slot application. 3. Safety training 1. Conduct radiation safety training in "Safety Training" in the BSRF user service system and take the exam. Only after passing the exam can you be added as a visitor. Safety training is valid for 2 years. 2. If the scores are not automatically synchronized successfully after passing the exam, please click "Synchronize Scores" in the upper right corner of the "Security Training" - "My Exams" page. 4. Application for visit 1. After the line station arranges the machine time, please visit the BSRF user service system to fill in the "Visit Application" at least 7 days before the start of the experiment and submit the information of the people who want to participate in the experiment. If you have any problems with the flight schedule, please communicate with the line station administrator in time. 2. Requirements for adding visitors: 1) Have a registered account on the hospital platform and fill in the ID number; 2) Pass the security examination and be within the validity period. 3. Visit time: The system defaults to the day before the experiment starts to the day after the experiment ends. 5. Registration before experiment 1. Users are required to bind the BSRF virtual access card before their first experimental visit (no need to bind it again for subsequent visits). Please complete the binding at least 3 days in advance according to the card number provided by the BSRF user service system email (internal users of the High Energy Institute do not need to bind). For binding instructions, please see the "Virtual Access Card Binding Operation Instructions for Users Outside the BSRF Institute". 2. Enter Hall 12 through the virtual access card (QR code), register before the experiment through the experimental self-service terminal and receive the dose card (see the "BSRF Experimental Self-Service Terminal Operation Instructions" for details). After about 5 minutes, activate the access control permissions for the corresponding experimental building and experimental hall. 6. Start the experiment Users use the virtual access card (QR code) and dosage card to enter the corresponding experimental hall to conduct experiments. 7. End of experiment 1. Return the dose card through the experimental self-service terminal (see the "BSRF Experimental Self-Service Terminal Operation Instructions" for details). Compensation is required if the dose card is damaged, lost or not returned. 2. Please check after the BSRF user service system sends the actual machine execution status. If you have any questions, please communicate with the line station administrator in time. 8. Experiment summary 1. Submit the experiment summary report through "Machine Time Reservation" - "My Machine Time" in the BSRF user service system within 2 months after the end of the experiment. 2. If there are any results (articles, patents, awards, social impact, etc.), please provide timely feedback in "My Results" in the BSRF user service system.
BSRF Beijing Synchrotron Radiation User Experiment Visit Process Instructions | # BSRF User Experiment Visit Process Instructions ## 1. Application for visit Please visit the BSRF user service system to fill in the visit application (https://user.bsrf.ihep.ac.cn/) at least 7 days before the start of the experiment. Visitors add requirements: 1. Have a registered account on the Chinese Academy of Sciences' Major Science and Technology Infrastructure Sharing Service Platform and fill in your ID number (https://lssf.cas.cn/); 2. Passed radiation safety training in BSRF user service system. ## 2. Bind all-in-one card account After receiving the access control card number sent by the BSRF user service system, the account binding is completed 3 days in advance. ## 3. Entering the park 1. Register and enter at the retractable gate of the park; 2. Enter Hall 12 through the access control virtual card (QR code). ## 4. Registration before experiment 1. Register before the experiment on the BSRF experiment self-service terminal placed in the foyer of Hall 12, including querying and selecting experiments, and receiving dose cards; 2. After about 5 minutes, the access control permissions for the experimental building, experimental hall and telescopic door corresponding to the selected experiment are activated. ## 5. Conduct experiments Scan the virtual access control card (QR code) against the access control QR code scanning head, and enter the corresponding line station to conduct the experiment. ## 6. Return after experiment After the experiment, return the dose card through the experimental self-service terminal (see the "BSRF Experimental Self-Service Terminal Operation Instructions" for details). About 5 minutes later, the access rights to the corresponding experimental building and experimental hall were cancelled, and the access rights to the door and telescopic door of Hall 12 were retained until 24:00 on the visit deadline.
Instructions for binding process of virtual access card for users outside BSRF Beijing Synchrotron Radiation Institute | Starting from May 2025, BSRF will enable virtual access cards (QR codes) for external users. The specific binding and usage instructions are as follows. 1. Involving access control and validity period instructions ● Telescopic door (Beijing Electron Positron Collider National Laboratory Park Door) Start time: About 5 minutes after the experiment self-service terminal completes pre-experiment registration End time: 24:00 of the visit end date in the BSRF user service system ● Hall 12 (Synchrotron Radiation East Hall) Door Start time: 00:00 of the visit start date in the BSRF user service system End time: 24:00 of the visit end date in the BSRF user service system ● Experimental hall door of Hall 12 (Synchrotron Radiation East Hall), Building door of Hall 15 (Synchrotron Radiation North Hall), Experimental hall door of Hall 15 (Synchrotron Radiation East Hall) Start time: About 5 minutes after the experiment self-service terminal completes pre-experiment registration End time: about 5 minutes after returning the dose card at the experimental self-service terminal 2. Preparation work Mobile phone number: Please use the mobile phone number provided when registering on the Chinese Academy of Sciences Major Science and Technology Infrastructure Sharing Service Platform All-in-one card number: Please check the all-in-one card number returned by the BSRF user service system via email (the email sent after the visit application is approved) Initial password: 123456 (initial password) 2. Binding steps Note: Please complete the binding at least three days before your visit. Step one: Follow the service account ● Open WeChat and search for the "Beijing Synchrotron Radiation User Window" service account ● Click "Follow" to enter the service account. Step 2: Log in and bind the access control virtual card. Step 2: Log in and bind the access control virtual card ● Click "User Center" → "Virtual Card" in the lower right corner of the service account ● Enter your mobile phone number → click "Next" ● Enter the card number and initial password 123456, and click "Bind" Step 3: Use access control virtual card (QR code) ● After successful login, find the "QR code" function on the virtual card interface ● Open the virtual access card in front of the access control equipment in each building and experimental hall ● Align the virtual access control card (QR code) with the QR code scanning head of the access control for verification## 4. Precautions ### If you encounter login problems, please confirm: - When authentication fails, please check whether the card number/password entered is correct (check the relevant email); - Whether to use the mobile phone number registered when applying. ### Other notes: - Please keep the screen brightness sufficient when using; - If the NFC function is turned on on the phone, the NFC card may be detected and the code cannot be scanned; - The virtual access card (QR code) will be automatically refreshed after 30 seconds of expiration, or after each use.
BSRF Beijing Synchrotron Radiation Institute User Access Card Authorization Instructions | # BSRF user access card authorization instructions Starting from May 2025, BSRF will enable one-card access authorization management for relevant buildings and experimental halls during the user experiment period. The details are as follows. 3. Involving access control and validity period instructions ● Telescopic door (Beijing Electron Positron Collider National Laboratory Park Door) Start time: About 5 minutes after the experiment self-service terminal completes pre-experiment registration End time: 24:00 of the visit end date in the BSRF user service system ● Hall 12 (Synchrotron Radiation East Hall) Door Start time: 00:00 of the visit start date in the BSRF user service system End time: 24:00 of the visit end date in the BSRF user service system ● Experimental hall door of Hall 12 (Synchrotron Radiation East Hall), Building door of Hall 15 (Synchrotron Radiation North Hall), Experimental hall door of Hall 15 (Synchrotron Radiation East Hall) Start time: About 5 minutes after the experiment self-service terminal completes pre-experiment registration End time: about 5 minutes after returning the dose card at the experimental self-service terminal 2. Authorization steps 1. Registration before experiment: Step 1: The day before the experiment, you can enter Hall 12 with your institute card; Step 2: Refer to the "BSRF Experiment Self-Service Terminal Operation Guide" and complete the pre-experiment registration operation on the experiment self-service terminal in Hall 12. About 5 minutes later, the one-card access control for the building and experimental hall corresponding to this experiment was opened. 2. Return after experiment: Return the dose card at the lab kiosk. The access rights to Hall 12 and the retractable door will be retained until the end date of the visit in the user service system, and the remaining access rights will be canceled approximately 5 minutes after the dose card is returned. Note: If the user card in the institute has access control permissions for the retractable door, Hall 12 and Hall 15, the access authorization during the experiment will not affect subsequent use.
HEPS-Biological macromolecule microcrystalline diffraction line station, official website | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262837.html
HEPS-Biological macromolecule microcrystalline diffraction line station, person in charge | High enhancement, gaozq@ihep.ac.cn
HEPS-biomacromolecule microcrystalline diffraction line station, PILOT contact person-1 | High enhancement, gaozq@ihep.ac.cn
HEPS-biomacromolecule microcrystalline diffraction line station, PILOT contact person-2 | She Zhun, shezhun@ihep.ac.cn
HEPS-biomacromolecule microcrystalline diffraction line station, scientific goal-1 | Efficiently and high-quality analysis of protein crystal structures with important biological functions
HEPS-biomacromolecule microcrystalline diffraction line station, scientific goal-2 | Establish a high-throughput, highly automated drug screening platform
HEPS-biomacromolecule microcrystalline diffraction line station, main parameters-energy range (keV) | 5–18 keV
HEPS-biomacromolecule microcrystalline diffraction line station, main parameter-energy resolution (ΔE/E) | ~2E-4
HEPS-biomacromolecule microcrystalline diffraction line station, main parameter-minimum spot size (H×V) | 1 µm × 1 µm
HEPS-biomacromolecule microcrystalline diffraction line station, main parameter-maximum spot size (H×V) | 18 µm × 28 µm
HEPS-biological macromolecule microcrystalline diffraction line station, main parameter-luminous flux | 1E13 phs/s
HEPS-biomacromolecule microcrystalline diffraction line station, sample environment | Low temperature, normal temperature
HEPS-biomacromolecule microcrystalline diffraction line station, detector | Single photon counting two-dimensional pixel array detector
HEPS-biomacromolecule microcrystalline diffraction line station, applied subject field | Life Sciences, Drug Development, Chemistry
HEPS-biomacromolecule microcrystalline diffraction line station, line station layout diagram-Picture 1 | [Picture placeholder]
HEPS-biomacromolecule microcrystalline diffraction line station, line station layout diagram-Picture 2 | [Picture placeholder]
HEPS-biomacromolecule microcrystalline diffraction line station, main features-title | High-brightness microfocus spot protein crystallography
HEPS-Biological macromolecule microcrystalline diffraction line station, main features-high brightness micro-focus spot protein crystallography-paragraph | This line station can provide adjustable light spot sizes from 1 µm×1 µm to 18 µm×28 µm. At the same time, the photon flux (phs/s) can reach 1E13 @12.4keV@200mA (1µm×1µm spot). In addition to satisfying routine protein crystallography studies, it also provides strong support for data collection of tiny protein crystals.
HEPS-biomacromolecule microcrystal diffraction line station, main features-diversified protein data acquisition methods-introduction | Based on the Arinax-MD2S diffractometer configured in this line station, this line station can provide a variety of crystal data collection strategies in addition to the traditional single-axis mode. For example:
HEPS - biomacromolecule microcrystalline diffraction line station, main feature - highly automated Raster Scan (grid scan) data collection - paragraph | In view of the difficulty in aligning the sample of microcrystals on the diffractometer and poor radiation resistance, the sample area is divided into grids with the same size as the light spot. A grid scanning method is used to quickly collect a diffraction pattern on each grid. The diffraction ability of each grid is obtained through preliminary processing and analysis of the diffraction pattern. Several grid points with strong diffraction ability are selected and 50 to 100 diffraction data are collected within a total rotation angle of 10 to 20 degrees. Then the diffraction data of multiple samples are automatically filtered and merged to obtain complete diffraction data.
HEPS-biomacromolecule microcrystalline diffraction line station, main features-helical step scanning experimental method-paragraph | If the crystal is larger than the beam size, the best practice to minimize radiation damage is to make the beam as large as the crystal. Spiral collection will help mitigate radiation damage by bringing a new volume of crystal into the beam for each image.
HEPS-biomacromolecule microcrystalline diffraction line station, main features-in-situ flat plate protein growth and collection mode-description | The data collection method of in situ flat plate protein crystals is a technology that directly collects X-ray diffraction data during the crystal growth process, avoiding the steps in traditional methods that require the crystals to be removed from the growth environment and frozen or transferred to other supports. This method is particularly suitable for studying the dynamic processes of protein crystals under natural growth conditions, such as crystal growth, phase transition or ligand binding.
HEPS-biomacromolecule microcrystal diffraction line station, main features-in-situ flat plate protein growth and collection mode-support | The diffractometer at this line station supports flat plate in-situ collection mode, which can study the structural characteristics of protein crystals in a room-temperature in-situ growth environment.
HEPS-biomacromolecule microcrystal diffraction line station, introduction to each part-main optical equipment-list | Horizontal reflective groove monochromator (HCCM, self-developed); compound refractive lens; fluorescence target, XBPM, slit/diaphragm
HEPS-biological macromolecule microcrystalline diffraction line station, introduction to each part-main experimental station equipment-list | High-precision diffractometer (ARINAX-MD2S); self-developed single-photon counting two-dimensional pixel array detector; automatic loading robot (Swordfish); liquid nitrogen cooling (Oxford Cryosystems 1000 series); experimental station control software (self-developed)
HEPS-biomacromolecule microcrystalline diffraction line station, line station equipment-sample environment | High performance diffractometer (ARINAX-MD2S)
HEPS-biomacromolecule microcrystalline diffractometer, line station equipment-high-performance diffractometer (ARINAX-MD2S)-paragraph | MD2S is a highly integrated diffractometer system, including air bearing rotating shaft, microscope system, sample illumination system, traditional goniometer head and replaceable flat plate dedicated goniometer head system, Beamstop, etc. Based on this diffractometer, multiple modes of protein crystallography diffraction experiments can be carried out. In addition to conventional crystal data collection, it also provides spiral step scanning mode, high-precision grid scanning mode, flat plate in-situ data collection mode, etc.
HEPS-biomacromolecule microcrystalline diffraction line station, line station equipment-domestic automatic loading manipulator (Swordfish)-paragraph | This line station uses the Swordfish, a loading manipulator independently developed by Shanghai Guangyuan. The loading manipulator can accommodate more than 20 pucks at the same time, and the sample changing time can be less than 20 seconds.
HEPS-biomacromolecule microcrystalline diffraction line station, line station equipment-liquid nitrogen cooling system (Oxford Cryosystems 1000 series)-paragraph | Using liquid nitrogen as the cooling medium can quickly and effectively reduce the sample temperature and provide a stable low-temperature environment, with an adjustable temperature of 80K-400K. Its precise temperature control function allows users to conduct experiments under different temperature conditions and ensures the stability and repeatability of experimental conditions. In addition, the Oxford Cryosystems 1000 Series features robust safety features to ensure the safety of operators and laboratory equipment. Due to its efficient and reliable performance, this series of liquid nitrogen cooling systems are widely used in various research and laboratory environments, promoting the progress of scientific research and the achievement of results.
HEPS-biological macromolecule microcrystalline diffraction line station, line station equipment-detector-paragraph | Self-developed single photon counting two-dimensional pixel array detector. The HEPS detector system team has conducted independent research on various key technologies of photon counting pixel detectors. Successfully developed a single-photon counting two-dimensional pixel array detector. The detector is based on the principle of over-threshold screening and realizes the front-end digitization of incident photons and high-speed readout through frame refresh. Compared with traditional CCD, CMOS cameras and other detectors, it has the characteristics of "zero noise" images, no dead time, and high refresh rate, and has become the mainstream detector in current synchrotron radiation light sources.
HEPS-biomacromolecule microcrystalline diffraction line station, line station equipment-software control system-paragraph | Based on PyQt, researchers at this line station developed a crystal diffractometer station control and data acquisition system dedicated to this line station to achieve complete control of the diffractometer MD2S. Realize the interface with EPICS, display the beam line status in real time, and control the movement of the diffractometer platform and detector platform. Integrate robot control to realize automatic loading of samples. And develop the Raster and helical data acquisition functions of microcrystals. It also implements advanced functions such as automatic heat map generation and positioning for grid scanning. In addition, this system can also achieve smooth communication and data exchange with the HEPS general software system (Mamba).
HEPS-biomacromolecule microcrystal diffraction line station, line station equipment-common accessories-list | Adjustable slit Aperture (10–150 µm variable); Capillary Capillary (reduce stray light); Beamstop (block direct beam and protect detector)
HEPS-biomacromolecule microcrystalline diffraction line station, line station equipment-supporting laboratory-paragraph | This line station provides a supporting laboratory with 4℃, -20℃ and -80℃ refrigerators to store samples, as well as ultra-clean benches, operating tables and other equipment.
HEPS-biomacromolecule microcrystalline diffraction line station, sample preparation-mode-1 | Traditional loop collection (liquid nitrogen freezing)
HEPS-biomacromolecule microcrystalline diffraction line station, sample preparation-mode-2 | In situ plate growth collection: plate specifications reference https://www.mitegen.com/product/crystaldirect-plate
HEPS-biomacromolecule microcrystalline diffraction line station, sample preparation-tip-1 | Since you need to switch devices in different modes, please be sure to confirm in advance which method you need to experiment.
HEPS-biomacromolecule microcrystalline diffraction line station, sample preparation-tip-2 | If you have special needs, please contact the line station staff in advance!
HEPS-biomacromolecule microcrystal diffraction line station, data collection and processing-collection software-paragraph | This station currently uses a self-developed software system for data collection. The system is convenient, efficient, and has diversified functions. It is constantly being optimized and developed. In the future, it will provide users with more diverse operations for data collection.
HEPS-biomacromolecule microcrystalline diffraction line station, data collection and processing-processing software-paragraph | The data processing software of this line station includes XDS, autoProc, xia2 and the domestic self-developed software autoPD.
HEPS-biomacromolecule microcrystal diffraction line station, data collection and processing-automation-paragraph | This site has initially implemented automated processing of data, using four software: xia2-xds, xia2-dials, autoProc and autoPD to automatically process the generated data in real time, allowing users to know the quality of the data immediately and help users make experimental decisions.
HEPS-Biological macromolecule microcrystalline diffraction line station, research field-description | Protein crystallographic structure analysis has wide applications in many fields. The following are examples of some important application fields:
HEPS-biomacromolecule microcrystalline diffraction line station, research field-drug design and development-description | ——
HEPS-biomacromolecule microcrystalline diffraction line station, research field-drug design and development-target protein structure analysis | Help design more effective drug molecules by analyzing the three-dimensional structure of drug target proteins.
HEPS-biomacromolecule microcrystalline diffraction line station, research field-drug design and development-drug-target interaction research | Reveal the binding mode of drug molecules and target proteins and optimize drug activity.
HEPS-Biological macromolecule microcrystalline diffraction line station, research field-drug design and development-research on drug resistance mechanism | Analyze the structure of mutant proteins, understand drug resistance mechanisms and design new drugs.
HEPS-Biological macromolecule microcrystalline diffraction line station, research field-enzymology and catalytic mechanism research-explanation | ——
HEPS-biomacromolecule microcrystal diffraction line station, research field-enzymology and catalytic mechanism research-enzyme active center analysis | Reveal the catalytic mechanism and substrate binding sites of enzymes.
HEPS-biomacromolecule microcrystal diffraction line station, research field-enzymology and catalytic mechanism research-enzyme engineering | Conduct directed evolution or rational design of enzymes based on structural information to improve enzyme activity or stability.
HEPS-biomacromolecule microcrystalline diffraction line station, research field-biomolecule interaction research-explanation | ——
HEPS-biomacromolecule microcrystalline diffraction line station, research field-biomolecular interaction research-protein-protein interaction | Analyze the structure of protein complexes and reveal the molecular mechanisms of signal transduction, metabolic regulation and other processes.
HEPS-biomacromolecule microcrystal diffraction line station, research field-biomolecular interaction research-protein-nucleic acid interaction | Study the interaction of transcription factors, ribosomes, etc. with DNA or RNA.
HEPS-Biological macromolecule microcrystalline diffraction line station, research field-structural biology and functional research-description | ——
HEPS-biomacromolecule microcrystalline diffraction line station, research field-structural biology and functional research-protein functional analysis | Infer protein functions through structural information, such as transporters, receptor proteins, etc.
HEPS-biomacromolecule microcrystalline diffraction line station, research field-structural biology and functional research-conformational change research | Reveal the conformational changes of proteins in different states, such as allosteric effects, folding and unfolding processes.
HEPS-Biological Macromolecule Microcrystalline Diffraction Line Station, Research Area-Pathogen Research and Vaccine Development-Description | ——
HEPS-Biological macromolecule microcrystalline diffraction line station, research field-pathogen research and vaccine development-viral protein structure analysis | Analyze the structure of key viral proteins (such as the spike protein) to provide a basis for vaccine design and antiviral drug development.
HEPS-Biological macromolecule microcrystalline diffraction line station, research areas-pathogen research and vaccine development-bacterial toxins and host interactions | Study the interaction mechanism between pathogen toxins and host proteins.
HEPS-biomacromolecule microcrystalline diffraction line station, research field-industrial biotechnology-description | ——
HEPS-biomacromolecule microcrystalline diffraction line station, research field-industrial biotechnology-industrial enzyme optimization | Optimize the performance of industrial enzymes (e.g., cellulases, lipases) through structural elucidation.
HEPS-biomacromolecule microcrystalline diffraction line station, research field-industrial biotechnology-biocatalysis | Designing efficient and specific biocatalysts.
HEPS-Biological macromolecule microcrystalline diffraction line station, research field-basic scientific research-description | ——
HEPS-Biological macromolecule microcrystalline diffraction line station, research field-basic scientific research-protein folding and stability | Study the folding mechanism and stability of proteins and reveal the relationship between protein misfolding and diseases.
HEPS-Biological macromolecule microcrystalline diffraction line station, research field-basic scientific research-evolution and functional relationship | By comparing the protein structures of different species, the evolution and functional relationships of proteins are studied.
HEPS-Biological Macromolecule Microcrystalline Diffraction Line Station, Research Area-Material Science and Nanotechnology-Description | ——
HEPS-Biological macromolecule microcrystal diffraction line station, research field-material science and nanotechnology-protein self-assembly | Study the self-assembly mechanism of proteins for the design of new biomaterials.
HEPS-Biological Macromolecule Microcrystalline Diffraction Line Station, Research Area-Materials Science and Nanotechnology-Nanostructure Design | Using protein structures to design functional nanomaterials.
HEPS-Biological Macromolecule Microcrystalline Diffraction Line Station, Research Area-Agriculture and Plant Science-Description | ——
HEPS-Biological Macromolecule Microcrystalline Diffraction Line Station, Research Area-Agriculture and Plant Science-Plant Disease Resistance Protein Research | Analyze the structure of plant disease-resistant proteins to help develop disease-resistant crops.
HEPS-Biological Macromolecule Microcrystal Diffraction Line Station, Research Area-Agriculture and Plant Science-Photosynthesis Mechanism | Study the structure of photosynthesis-related proteins and optimize plant photosynthetic efficiency.
HEPS-biomacromolecule microcrystalline diffraction line station, research field-medical diagnosis and treatment-description | ——
HEPS-biomacromolecule microcrystalline diffraction line station, research field-medical diagnosis and treatment-disease-related protein research | Analyze the structure of proteins related to cancer, neurodegenerative diseases, etc., and develop diagnostic markers or therapeutic targets.
HEPS-biomacromolecule microcrystalline diffraction line station, research field-medical diagnosis and treatment-antibody engineering | Design efficient antibodies based on structural information for disease treatment or diagnosis.
HEPS-Biological macromolecule microcrystalline diffraction line station, further reading-application for the first round of trial project (PILOT proposal) | (Please visit the extended reading link on the official website page)
HEPS-Engineering Materials Line Station, official website page | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262778.html
HEPS-Engineering Materials Line Station, Title | Engineering Materials Line Station
HEPS-Engineering Materials Line Station, Introduction | The Engineering Materials Line Station is an insert line station dedicated to studying the physical and chemical properties of materials under real working conditions. Utilizing high-brightness, high-energy X-rays and integrating a variety of on-site in-situ sample environments can provide multi-scale, rapid and non-destructive characterization for materials research throughout the life cycle. The Engineering Materials Line Station contains two independent experimental sheds, one dedicated to research in materials engineering and materials chemistry.
HEPS - Engineering Materials Line Station, Materials Engineering Experiment Shed - Supported Experimental Methods | Diffraction (single crystal/powder diffraction, depth/time resolved diffraction, XRD-CT), imaging (full field/scan+absorption/phase contrast+projection imaging/CT, grain-resolved diffraction contrast imaging), scattering (PDF)
HEPS - Engineering Materials Line Station, Materials Engineering Experiment Shed - Energy Range | 50–170 keV
HEPS - Engineering Materials Line Station, Materials Engineering Experiment Shed - Energy Resolution | 10^{-3}
HEPS-engineering material line station, materials engineering experimental shed-sample spot size | ~1 mm × 1 mm (maximum), 5 μm × 5 μm (minimum)
HEPS-Engineering Materials Line Station, Materials Engineering Experiment Shed-Light Flux at Sample | 2×10^{12} phs/s (@100 keV, 200 mA, unfocused)
HEPS-engineering material line station, materials engineering experimental shed-sample stage load | >200 kg (comprehensive), >1.5 t (heavy load)
HEPS - Engineering Materials Line Station, Materials Chemistry Experiment Shed - Supported Experimental Methods | Scattering (SAXS, WAXS, SAXS+WAXS, time-resolved SAXS, PDF), diffraction (single crystal/powder diffraction), imaging (full field/scan + absorption/phase contrast + projection imaging/CT)
HEPS-Engineering Materials Line Station, Materials Chemistry Experiment Shed-Energy Range | 50–170 keV
HEPS - Engineering Materials Line Station, Materials Chemistry Experiment Shed - Energy Resolution | 10^{-3}
HEPS-engineering material line station, materials chemistry experimental shed-spot size at the sample | ~0.7 mm × 0.7 mm
HEPS-Engineering Materials Line Station, Materials Chemistry Experiment Shed-Light Flux at Sample | 2×10^{12} phs/s (@100 keV, 200 mA, unfocused)
HEPS-Engineering Materials Line Station, Materials Chemistry Experiment Shed-Sample Stage Load | >200 kg
HEPS-Engineering Materials Line Station, Line Station Person in Charge | Yang Yiming yangym@ihep.ac.cn
HEPS-engineering material line station, optical path layout diagram | [Picture placeholder] Light path layout diagram of engineering materials beamline station
HEPS-pink small angle scattering station, official website | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262838.html
HEPS-Pink Small Angle Scattered Ray Station, title | Pink light small angle scattering ray station
HEPS-pink light small angle scattering station, person in charge | Mo Guang, mog@ihep.ac.cn
HEPS-pink light small angle scattered line station, PILOT contact person | Mo Guang, mog@ihep.ac.cn
HEPS-pink light small angle scattering station, scientific goal | The Pink Light Small Angle Scattering Station will provide users with experimental tests related to X-ray small angle scattering (SAXS). SAXS technology provides comprehensive structural information such as nanoscale particle size, shape, size distribution, periodicity, and fractal within the sample by measuring the scattering intensity distribution of X-rays in a small angle range after passing through the sample. Because SAXS technology adopts exposure-based acquisition, the acquisition time can be reduced to the order of milliseconds or even microseconds as the light flux increases. Therefore, it is suitable for dynamic changes in the structural characteristics of materials at the nanoscale, and is called an indispensable analysis tool in the fields of nanoscience, structural biology, and soft matter science. In addition, SAXS has a wide application range for samples and is simple to prepare. The sample state can be liquid, powder, film or even a mixture of them. SAXS tests are generally non-destructive and can be used repeatedly or for other measurements. In order to meet the regular testing needs of the majority of users, the Pink Light Small Angle Scattering System pursues the construction of a multi-functional SAXS line station with simple structure, high stability, high throughput, good statistics. Overview of the overall system design: a thin beam of fundamental wave radiation (8–12 keV) is generated by a vacuum external undulator with an energy bandwidth of 1.5% (this quasi-monochromatic light is called pink light). It is horizontally deflected and collimated and irradiated onto the sample. The scattered signals are received by the SAXS, WAXS and USAXS detectors placed in the optical path, respectively, to obtain scattered signals in different angle ranges.
HEPS-pink light small angle scattering station, energy range keV | 8–12 keV
HEPS-pink light small angle scattering station, energy resolution ΔE/E | 1%
HEPS-pink light small angle scattering line station, minimum spot H×V | 100 µm × 100 µm
HEPS-pink light small angle scattering station, maximum spot H×V | 500 µm × 500 µm
HEPS-Pink light small angle scattering ray station, luminous flux | 5E14
HEPS-pink light small angle scattering station, sample environment | High temperature, low temperature, stretching, liquid phase synthesis
HEPS-pink light small angle scattering station, detector | Eiger 4M
HEPS-pink light small angle scattering station, applied subject area | chemistry, materials, physics
HEPS - pink light small angle scattering station, main feature - high photon flux | About 10^{15} cps, suitable for weak scattering signal and dynamic scattering detection.
HEPS - pink light small angle scattering station, main features - large measurement range | Å–nm–μm, meeting the needs of studying samples with rich structural dimensions.
HEPS - pink light small angle scattering station, main features - multiple focus modes | A variety of focusing modes with focused light spots at different positions are available.
HEPS-pink light small angle scattering station, main feature-point circular light spot | Suitable for the measurement of anisotropic samples.
HEPS-pink light small angle scattering station, main features-detector combination | Vacuum version of photon counting detector with high data quality.
HEPS-pink light small angle scattering station, main experimental method-conventional transmission type | WAXS, SAXS, USAXS technology or combined measurements.
HEPS-pink light small angle scattering line station, main experimental method-grazing incidence measurement | Measurements of GiWAXS, GiSAXS technologies.
HEPS-pink light small angle scattering station, main experimental method-time resolution | For fast changing processes, dynamic scattering detection with millisecond time resolution (T-SAXS) can be provided.
HEPS-pink light small angle scattering station, introduction to each part-insert | The insert is a traditional classic external vacuum undulator IAU25 with a fundamental wave energy of 8–12 keV.
HEPS-pink light small angle scattering beam station, introduction to each part-beam line | In order to maintain the advantage of high throughput, some experiments can directly use the quasi-monochromatic light generated by the undulator as the incident light. For experiments with higher energy resolution (angular resolution) requirements, this can be compensated by subsequent addition of moveable monochromators.
HEPS-pink light small angle scattering station, introduction to each part-experimental station | The camera pipeline is 23 m long and equipped with 3 Eiger detectors for atomic size (wide angle), nanometer size (small angle) and micron scale measurements.
HEPS-pink light small angle scattering line station, line station equipment-sample environment | Xiaojiao Station is currently equipped with self-made high-temperature environment, low-temperature environment, two-way stretching machine, four-way stretching machine, in-situ liquid phase nucleation growth and other sample environments.
HEPS-pink light small angle scattering line station, line station equipment-detector-WAXS | Model: Eiger 4M; detection angle: 5–50; pixel size: 75 µm × 75 µm; frequency: 100 Hz.
HEPS-pink light small angle scattering line station, line station equipment-detector-SAXS-1 | Model: Eiger 4M; detection angle: 0.04–6; pixel size: 75 µm × 75 µm; frequency: 100 Hz.
HEPS-pink light small angle scattering line station, line station equipment-detector-SAXS-2 | Model: Eiger 1M; detection angle: 0.004–0.4; pixel size: 75 µm × 75 µm; frequency: 2000 Hz.
HEPS - pink light small angle scattering line station, line station equipment - sample preparation equipment | Xiaojiao Station is equipped with centrifuges, cooling water circulation machines, ball mills, pipettes and other sample preparation equipment.
HEPS-pink light small angle scattering station, data collection and processing | ID08 line station provides small angle scattering data processing software Daisy-SAXS. Daisy-SAXS is a data processing software for small-angle scattering experiments written in Python. It provides image import, preview, preprocessing, spot center calibration, mask generation, azimuth angle integration, similarity judgment, guinier fitting, BIFT fitting, molecule width calculation, result visualization and other modules. It supports data processing in formats such as tiff and HDF5. Daisy-SAXS provides a Web front-end user interface based on the Jupyter ecosystem and is deployed on the HEPS interactive computing Torch platform. Users can use Daisy-SAXS to access experimental data through the deployed thin terminal on the experimental station, and call HEPS computing resources for small-angle scattering data processing.
HEPS-pink light small angle scattering station, examples of application fields | SAXS is a technology and method for studying the microstructure and morphological characteristics of substances. The objects studied are structures much larger than the atomic size, ranging from nanoscale to micron scale, so it involves a wide range of samples. Generally speaking, SAXS research objects mainly include nanomaterials, mesoporous materials, biological macromolecules, polymers, etc.
HEPS-pink light small angle scattering station, further reading | Application for the first round of trial project (PILOT proposal)
HEPS-pink light small angle scattering station, related links | Chinese Academy of Sciences; Chinese Academy of Sciences Major Infrastructure Sharing Platform; Beijing Synchrotron Radiation Facility; Shanghai Light Source; Hefei Light Source; Information Monitoring Platform of the Institute of High Energy - SR Device Updates.
HEPS-Pink Small Angle Scattering Station, contact information | Address: HEPS Engineering Office, No. 1 Guangyuan Street, Huairou District, Beijing; Postal code: 101407; Contact person: Yuan Mengyao (heps@ihep.ac.cn); HEPS user office: Yu Meijuan (heps-user@ihep.ac.cn).
HEPS-Pink light small angle scattering line station, line station pictures | [Picture placeholder]
HEPS-low dimensional structure probe line station, official website | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262836.html
HEPS-Low Dimensional Structure Probe Line Station, person in charge | Wang Huanhua (wanghh@ihep.ac.cn)
HEPS-Low Dimensional Structure Probe Line Station, PILOT Contact-1 | Wang Huanhua (wanghh@ihep.ac.cn)
HEPS-Low Dimensional Structure Probe Line Station, PILOT Contact-2 | Weng Xiaorong (wengxr@ihep.ac.cn)
HEPS-Low Dimensional Structural Probe Line Station, Title | Low Dimensional Structural Probe Line Station (LoDiSP)
HEPS-Low Dimensional Structure Probe Line Station, Scientific Target | Using surface X-ray scattering (SXRS) as the core method, the in-situ structure and dynamic changes of low-dimensional materials and devices in various external field environments are analyzed.
HEPS-low-dimensional structure probe line station, line station design | An In Vacuum Undulator (IVU) insert is used as the light source; a beam line connects two experimental stations in series, and each experimental station contains two experimental areas (2+2 structure).
HEPS-low dimensional structure probe line station, EH1 introduction | The former experimental station (EH1) carries out X-ray scattering experiments in multiple environments. The core equipment is a 6+3 circular diffractometer, equipped with a variety of small sample environments. It provides resonant and non-resonant surface diffraction (SXRD), reflection (XRR), crystal truncated rod (CTR) analysis, reciprocal space mapping (RSM), surface atomic pair distribution function (SPDF), photon correlation spectroscopy (XPCS) and other methods.
HEPS-low dimensional structure probe line station, EH2 introduction | The rear experimental station (EH2) performs surface scattering in-situ structural analysis under ultra-high vacuum. The core is the "SXRS + HMBE + RHEED + XPCS + XRR" in-situ structural probe integrated device. The hybrid MBE cavity can be pushed away from the diffractometer and replaced with the user's own in-situ sample preparation cavity.
HEPS-low dimensional structure probe line station, scientific goal-refinement | (1) Study the structure and dynamic changes of low-dimensional materials and devices in an external field environment; (2) Study the relationship between the structure and physical properties of various solid interfaces and solid-liquid interfaces.
HEPS-low-dimensional structure probe line station, main features-incident light degree of freedom | The energy is adjustable, the polarization is adjustable, and it supports downward deflection.
HEPS-low-dimensional structure probe line station, main feature-highly coherent beam | The beam is highly coherent, allowing experiments such as GIXPCS, WAXPCS, and BCDI to be carried out.
HEPS - low-dimensional structure probe line station, main features - diverse sample environments | It can provide a variety of environments such as in-situ growth, temperature, electric field, magnetic field, stress, electrochemical reaction, etc.
HEPS - low-dimensional structure probe line station, main feature - directional resonance excitation | Combining polarization and energy tunability to achieve directional resonance excitation.
HEPS-low-dimensional structure probe line station, main features-element identification | It has resonance scattering and fluorescence spectroscopy capabilities to achieve element identification.
HEPS-low-dimensional structure probe line station, main features-dynamic detection | It has dynamic detection capabilities of time-resolved X-ray scattering and GI-/WA-XPCS.
HEPS-low-dimensional structure probe line station, main features-polarization analysis | Scattered light polarization analysis is realized based on 6+3 circular diffractometer.
HEPS-low-dimensional structure probe line station, main features-low-dimensional structure analysis | Low-dimensional structural analysis of ultrathin films, surfaces and interfaces.
HEPS-low-dimensional structure probe line station, main features-fast CTR analysis | Relying on high-energy light source and micro-focusing conditions to achieve rapid CTR analysis.
HEPS-low-dimensional structure probe line station, main features-EH1 integration | EH1 uses a retractable flight channel, integrating WA-XPCS and SAXPCS, as well as WAXRD and GIXRD.
HEPS-low-dimensional structure probe line station, main features-EH2 integration | EH2 will build a "hybrid MBE + XPCS integrated" in-situ structural probe; it can also build an in-situ structural probe integrated device that uses evanescent waves to modulate the film growth morphology.
HEPS-low-dimensional structure probe line station, main equipment | Huber 6+3 circle diffractometer; flight tube (retractable).
HEPS-low-dimensional structure probe line station, sample environment-high temperature | 300–1573 K.
HEPS-low-dimensional structure probe line station, sample environment-low temperature | 5–300 K.
HEPS-low-dimensional structure probe line station, sample environment-vacuum | 10^{-7} Pa.
HEPS-low-dimensional structure probe line station, detector-area detector | Eiger 4M.
HEPS-Low Dimensional Structure Probe Line Station, Detector-Point Detector | 9910.
HEPS-low-dimensional structure probe line station, main experimental methods | XRD, GIWAXS, PDF, XPCS, etc.
HEPS-Low Dimensional Structural Probe Line Station, Daisy-PDF | Full scattering PDF data processing software based on Python, supports high-performance parallel batch processing; includes image import, preview, preprocessing, azimuth angle integration, PDF function calculation and visualization modules; supports tiff, HDF5 and other formats; provides a Jupyter ecological web front-end, deployed on the HEPS interactive computing Torch platform, and can access experimental data through thin terminals and call HEPS computing resources for PDF processing.
HEPS-low dimensional structure probe line station, Daisy-XPCS | X-ray photon correlation spectrum software based on Python, including image import, mask/ROI selection, preprocessing, g2 and c2 calculation, visualization and fitting analysis, etc.; supports online/offline reconstruction; supports tiff and HDF5 formats; supports CPU/GPU acceleration; provides Jupyter ecological Web front-end, deployed on the HEPS Torch platform, users can access data through thin terminals and call HEPS GPU resources for XPCS processing.
HEPS-low-dimensional structure probe line station, line station picture-1 | [Picture placeholder]
HEPS-low-dimensional structure probe line station, line station picture-2 | [Picture placeholder]
HEPS-hard X-ray coherent scattering station, official website | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262829.html
HEPS - Hard X-ray Coherent Scattering Station, Introduction | As a dedicated coherent line station on HEPS, the hard X-ray coherent scattering station will make full use of the high brightness and high coherence of HEPS. It mainly includes two major types of coherent experimental methods, coherent X-ray diffraction imaging (CDI) and X-ray photon correlation spectroscopy (XPCS).
HEPS - Hard X-ray coherent scattering station, CDI (Coherent Diffraction Imaging) | For CDI, it specifically includes scanning coherence diffraction imaging (Ptychography), Bragg CDI (BCDI) and SAXS The three experimental technologies of CDI are mainly used for nanometer-resolution 3D structure imaging and strain imaging of samples. It can also be combined with X-ray fluorescence and absorption spectroscopy to achieve element imaging and valence state imaging. For example, it can be used for high-resolution 3D imaging of cells, providing the spatial arrangement relationship of various organelles in the cell and the fine structure of the organelles. It can be used to determine the 3D structure of metal glass shear bands, thereby providing the physical mechanism of plastic deformation and yielding. It can be used for various nanocrystal strains. For example, it helps to establish the relationship between catalytic activity and strain in catalysts, in-situ measurement of the defect distribution and movement of grains in lithium-ion battery cathode materials in battery materials, and measurement of the boundaries and defects of individual grains in polycrystalline materials. It can also be used to determine the structure of various micro-nano electronic devices such as integrated circuits. High-resolution 3D imaging of integrated circuits can help understand their functions, provide direct feedback on design and processing processes, quality control, and reverse engineering.
HEPS - Hard X-ray Coherent Scattering Station, XPCS (Photon Correlation Spectroscopy) | For XPCS, including small-angle and wide-angle modes, they can obtain the basic dynamic functions of the system, such as intermediate scattering functions and dynamic structure factors, etc., which are mainly used for the measurement of equilibrium dynamics, non-equilibrium dynamics and dynamic heterogeneity of condensed matter; for example, XPCS can be used to study the relaxation dynamics of various glass systems and colloidal systems. During the glass transition process, XPCS can completely measure α relaxation and β relaxation in glass transition, which helps to improve the understanding of glass transition and the nature of glass; it can be used to measure the fluctuations of various order parameters in strongly correlated electronic systems, including charge order, spin order, magnetic domain, etc.; it can also be used to measure surface interface dynamics, such as diffusion behavior, self-assembly behavior and kinetic transitions at the nanoparticle emulsion interface.
HEPS-hard X-ray coherent scattering line station, line station design index-energy range | 7–25 keV
HEPS-hard X-ray coherent scattering line station, line station design index-energy resolution | 10^{-4}
HEPS-hard X-ray coherent scattering line station, line station design index-coherent flux | Greater than 10^{12} ph/s; adjustable coherence (coherence component 0.5–0.95)
HEPS-hard X-ray coherent scattering line station, line station design index-Ptychography | Spot size 100 nm–5 μm; resolution 1–50 nm; field of view 1–100 μm.
HEPS-hard X-ray coherent scattering line station, line station design index-SAXS CDI | Spot size 2–20 μm; resolution 1–10 nm; field of view 1–10 μm.
HEPS-hard X-ray coherent scattering line station, line station design index-BCDI | Spot size 100–500 nm; resolution 1–10 nm; field of view 50–500 nn (sic).
HEPS-hard X-ray coherent scattering line station, line station design index-SAXS XPCS | Spot size 10–20 μm; spatial scale 1–200 nm; time scale 1000 s–1 μs.
HEPS-hard X-ray coherent scattering line station, line station design index-WAXS XPCS | The spot size is 100 nm–2 μm; the spatial scale is 0.1–1 nm; the time scale is 1000 s–1 μs.
HEPS-Hard X-ray Coherent Scattering Station, person in charge | Zhou Liang (zhouliang@ihep.ac.cn)
HEPS-hard X-ray coherent scattering station, top view of optical structure | [Picture placeholder]
HEPS-hard X-ray coherent scattering station, schematic diagram of the experimental station | [Picture placeholder]
HEPS-hard X-ray nanoprobe line station, official website page | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262781.html
HEPS-hard X-ray nanoprobe line station, title | Hard X-ray Nanoprobe Line Station
HEPS - Hard X-ray Nanoprobe Line Station, Introduction | The hard X-ray nanoprobe line station is one of the three ultra-long beamline stations of the HEPS light source. Its biggest feature is that it uses the extremely low emittance characteristics of the HEPS light source to focus hard X-rays to the nanometer scale to form an ultra-high brightness optical probe, which is used to carry out high-resolution in-situ bulk phase characterization research. Combining multi-modal characterization with various experimental techniques such as X-ray imaging, diffraction, spectroscopy, and electron microscopy imaging, key information such as two-dimensional or three-dimensional structure, strain, chemical state, and dynamic effects of materials can be analyzed at the nanoscale or even atomic scale. This line station can be widely used in cutting-edge scientific research in the fields of materials, life, physics, chemistry, environment, semiconductors, medicine, cultural heritage and planetary science.
HEPS-hard X-ray nanoprobe line station, main experimental method-nXRF | High-resolution fluorescence imaging (n-XRF).
HEPS-hard X-ray nanoprobe line station, main experimental method-Ptychography | Scanning coherence diffraction imaging (Ptychography).
HEPS-hard X-ray nanoprobe line station, main experimental method-Nano-holography | Nano-holographic coherence imaging (Nano-holography).
HEPS-hard X-ray nanoprobe line station, main experimental method-nXRD | Nanodiffraction (n-XRD).
HEPS-hard X-ray nanoprobe line station, main experimental method-nXANES | Near-edge absorption spectroscopy (n-XANES).
HEPS-hard X-ray nanoprobe line station, main experimental method-SEM | Online scanning electron microscopy (SEM).
HEPS-hard X-ray nanoprobe line station, main parameters-energy range | 5–30 keV
HEPS-hard X-ray nanoprobe line station, main parameters-energy resolution | 10^{-4}~10^{-5}
HEPS-hard X-ray nanoprobe line station, main parameters-minimum focus spot-high resolution mode | <10 nm
HEPS-hard X-ray nanoprobe line station, main parameters-minimum focus spot-in situ mode | <50 nm
HEPS-hard X-ray nanoprobe line station, main parameters-luminous flux-high resolution mode | 10^{9}~10^{10} phs/s @10 keV
HEPS-hard X-ray nanoprobe line station, main parameters-luminous flux-in situ mode | 10^{10}~10^{11} phs/s @10 keV
HEPS-hard X-ray nanoprobe line station, main parameters-focused spot working distance-high resolution mode | 1–2 mm
HEPS-hard X-ray nanoprobe line station, main parameters-focused spot working distance-in-situ mode | 40–50 mm
HEPS-hard X-ray nanoprobe line station, main parameters-sample size | Sub-micron to millimeter scale (sub-μm ~ mm)
HEPS-hard X-ray nanoprobe line station, main parameters-imaging resolution | Sample information (structure, elements, valence state, etc.) imaging resolution is approximately nanometer scale
HEPS-Hard X-ray Nanoprobe Line Station, Person in Charge | Chang Guangcai (changgc@ihep.ac.cn)
HEPS-hard X-ray nanoprobe line station, optical layout top view | [Picture placeholder] Top view of the optical layout of the hard X-ray nanoprobe line station
HEPS-hard X-ray imaging line station, official website page | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262832.html
HEPS-Hard X-ray Imaging Line Station, Title | Hard X-ray Imaging Line Station
HEPS-Hard X-ray Imaging Line Station, Introduction | The hard X-ray imaging line station is an insert line station, which has the advantages of high sensitivity, strong penetration, multiple spatial and temporal resolutions, large field of view, multi-contrast mechanism fusion, and multi-sample environment compatibility. It simultaneously achieves the advantages that ordinary line stations cannot have, such as large field of view and high resolution, strong penetration and high sensitivity. The experimental methods are mainly hard X-ray coaxial phase contrast imaging and diffraction contrast imaging; application fields include engineering materials and components, biomedicine, paleontology, etc.
HEPS - Hard X-ray Imaging Line Station, Insert Light Source and Energy Spectral Range | Three insert light sources, CPMU, IAW and Mango Twister, are used for switching. The pink light spectrum range is 5keV-300keV and the monochromatic light spectrum range is 40keV-200keV. It provides the best X-ray energy for non-destructive 3D imaging of various samples. The maximum source distance is 330m. SRX has high spatial coherence, small point spread function of the light source to the sample, and large light spot at the sample. The lateral size of the light spot is variable from 0.01mm to 300mm, and the spatial resolution is variable from 0.5μm to 50μm. It uses high-stability and wide-bandwidth monochromators such as optical filters, compound refractive lenses, and bendable adjustable Laue twin crystal monochromators to provide multi-scale light spots with high stability, high throughput, high uniformity, high coherence, and wide energy spectrum range.
HEPS - Hard X-ray Imaging Line Station, Detector | We independently design and develop a series of high-performance hard
HEPS-hard X-ray imaging line station, data management and calculation | Equipped with software and hardware systems for the management, transmission, reconstruction and analysis of 250TB of high data flow per day.
HEPS-Hard X-ray Imaging Line Station, person in charge | Li Gang (lig@ihep.ac.cn)
HEPS-Hard X-ray Imaging Line Station, Line Station Picture-1 | [Picture placeholder]
HEPS-Hard X-ray Imaging Line Station, Line Station Picture-2 | [Picture placeholder]
HEPS-Structural Dynamics Line Station, official website page | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262828.html
HEPS - Structural Dynamics Line Station, Title | Structural Dynamics Line Station
HEPS-Structural Dynamics Line Station, general introduction | The Structural Dynamics Line Station is an insert line station that focuses on ultra-fast time-resolved in-situ, real-time dynamic characterization of irreversible processes such as additive manufacturing and dynamic loading. On the basis of synchrotron radiation phase contrast imaging (PCI), diffraction and scattering experimental methods, combined with ultrafast time resolution technology, the dynamic deformation, microscopic defects and organizational evolution of advanced materials during the entire cycle from manufacturing to service are studied. The line station uses two short-period CPMUs to ensure ultra-high single-shot intensity in the high-energy range and achieve a multi-scale characterization method with a time resolution of up to 420 ps and a spatio-temporal resolution of up to 50 nm.
HEPS - Structural Dynamics Line Station, subject area | solid state physics
HEPS - Structural Dynamics Line Station, subject area | Engineering materials
HEPS - Structural Dynamics Line Station, subject area | planetary science
HEPS - Structural Dynamics Line Station, subject area | Mechanics of Materials
HEPS - Structural Dynamics Line Station, Scientific Applications - High Energy Density Physics | Research on phase transition under extreme conditions of high-power laser-driven materials and high-pressure equation of state.
HEPS - Structural Dynamics Line Station, Scientific Application - Dynamic Loading of Advanced Materials | Material deformation, fatigue and failure behavior under full strain rate loading conditions provided by air cannons, Hopkins rods, in-situ tensile fatigue and other devices.
HEPS-Structural Dynamics Line Station, Scientific Application-In-situ Real-time Characterization of the Additive Manufacturing Process | The generation and evolution rules of microstructure and defects during laser and electron beam additive manufacturing.
HEPS - Structural Dynamics Line Station, Experimental Methods | Time-resolved X-ray diffraction
HEPS - Structural Dynamics Line Station, Experimental Methods | Time-resolved X-ray phase contrast imaging
HEPS - Structural Dynamics Line Station, Experimental Methods | Time-resolved X-ray scattering
HEPS - Structural Dynamics Line Station, Experimental Methods | Time-resolved X-ray nanoimaging
HEPS - Structural Dynamics Line Station, Optical Mode - White Light Mode | Spot size 1 mm–14 mm; single pulse intensity >10^{9} phs/pulse; luminous flux >10^{16} phs/s.
HEPS - Structural Dynamics Line Station, Optical Mode - Micron Mode | Spot size 5 μm–500 μm; luminous flux >10^{15} phs/s; photon energy 20–70 keV; energy resolution 0.5%–2.8%.
HEPS - Structural Dynamics Line Station, Optical Mode - Nanomode | Spot size <50 nm; luminous flux >10^{14} phs/s; photon energy 21.8 keV; energy resolution 1.2%.
HEPS-structural dynamics line station, in-situ device-dynamic loading device | High-power laser (100 J), first-level air cannon, second-level air cannon (up to about 18 m), Hopkinson rod (pull and compression rod).
HEPS-Structural Dynamics Line Station, In-situ Device-Additive Manufacturing Device | Laser powder feeding, laser powder laying, and electron beam additive manufacturing.
HEPS - Structural Dynamics Line Station, Detector - Indirect Imaging Detector | Ultra-high-speed cameras (KIRANA, Specialized Imaging; Fastcam SA‑Z, Photron).
HEPS - Structural Dynamics Line Station, Detector - Direct Scattering Detector | High-energy pixel array detectors (Keck‑PAD‑CdTe, Sydor; Eiger‑CdTe, Dectris).
HEPS-Structural Dynamics Line Station, person in charge | Zhang Bingbing (010-88235978, zhangbb@ihep.ac.cn)
HEPS-Structural Dynamics Line Station, Schematic of Beamline Optical Element Layout | [Picture placeholder]
HEPS-structural dynamics line station, schematic diagram of equipment layout in experimental shed | [Picture placeholder]
HEPS-X-ray microscopy imaging line station, official website page | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262842.html
HEPS-X-ray microscopy line station, title | X-ray microscopy imaging line station
HEPS-X-ray microscopy imaging line station, person in charge/PILOT contact person | Yuan Qingxi (yuanqx@ihep.ac.cn)
HEPS-X-ray microscopy imaging line station, line station review-scientific goals | Nano-resolution full-field imaging based on zone plate amplification can quickly obtain high-resolution three-dimensional structural information, three-dimensional element distribution information and three-dimensional element valence distribution information of the sample.
HEPS-X-ray microscopy imaging line station, line station overview-main parameters-highest spatial resolution of imaging | 20 nm
HEPS-X-ray microscopy imaging line station, line station overview-main parameters-energy range | 5–15 keV
HEPS-X-ray microscopy imaging line station, line station overview-main parameters-photon flux at the sample | ~1E11 phs/s
HEPS-X-ray microscopy imaging line station, line station overview-main parameters-energy resolution (ΔE/E) | ≈1E-4
HEPS-X-ray microscopy imaging line station, line station overview-line station layout diagram | [Picture placeholder]
HEPS-X-ray microscopy imaging line station, line station overview-main features | Full-field imaging as low as 20 nm obtains high-resolution three-dimensional structural information; element absorption edge subtraction imaging obtains high-resolution three-dimensional element distribution information; XANES imaging obtains high-resolution three-dimensional distribution information of element valence states.
HEPS-X-ray microscopy imaging line station, line station equipment-sample environment | (The page section exists, no specific entries are listed)
HEPS-X-ray microscopy imaging line station, line station equipment-detector-1 | Hamamatsu ORCA-Flash4.0
HEPS-X-ray microscopy imaging line station, line station equipment-detector-2 | Hamamatsu ORCA-Quest2
HEPS-X-ray microscopy imaging line station, line station equipment-detector-3 | Andor ZYLA5.5X-FO
HEPS-X-ray microscopy imaging line station, line station equipment-sample preparation equipment-1 | Femtosecond laser processing equipment
HEPS-X-ray microscopy imaging line station, line station equipment-sample preparation equipment-2 | Sample pre-alignment system: measuring microscope + four-axis micro-operating system
HEPS-X-ray microscopy imaging line station, line station equipment-common accessories-1 | Spotting capillary
HEPS-X-ray microscopy imaging line station, line station equipment-common accessories-2 | silicon nitride window
HEPS-X-ray microscopy imaging line station, line station equipment-supporting laboratory | HEPS assisted laboratory
HEPS-X-ray microscopy imaging line station, data collection and processing - data collection | The TXM line station uses the high-energy synchrotron radiation light source software system and the data acquisition software MambaBE jointly designed and developed by the ID30 line station for data collection; image acquisition modes include: Single / Count / Continuous; Average / Focal series / 2D mosaic; Tomography (stepping) / Fly CT; Tomo + bg (user-defined background scheme); Mosaic CT (stepping and flying); 2D XANES / 3D XANES.
HEPS-X-ray microscopy imaging line station, data collection and processing-data processing-Mocupy | Mocupy: A stand-alone CT reconstruction software developed by HEPS ID30 line station, which realizes GPU-accelerated CT rapid reconstruction based on CUDA. It includes modules such as image preprocessing, jitter correction, phase recovery, CT reconstruction and data analysis, and can realize rapid reconstruction of CT data offline.
HEPS-X-ray microscopy imaging line station, data collection and processing-data processing-WebCT | WebCT: The web version of HEPS CT reconstruction software, which is deployed online by the ID30 line station and the computing and network communication system, and uses HEPS computing power to achieve online image preprocessing and CT reconstruction.
HEPS-X-ray microscopy imaging line station, research highlight-Nano-resolution spectroscopy CT (lithium-rich NMC) | TXM near-edge spectroscopy imaging was used to obtain 3D XANES data of the initial state and 4.8V charged state particles of the lithium-rich NMC battery material (LirNMC). It was found that: the initial state Mn is in a high valence state and is relatively uniformly distributed; in the charged state, Mn has a distribution of "higher on the surface and middle, lower in the middle layer"; the 3D XANES results of Co and Ni also show significant inhomogeneity and are related to depth. The results show that the structural and chemical complexity of battery cathode materials at the mesoscopic scale is expected to help understand and solve the voltage attenuation problem of lithium-rich cathodes (Jin Zhang et al., Nat Commun, 2020, 11:6342).
HEPS-X-ray microscopy imaging line station, resource download | CT reconstruction software Mocupy, https://ihep.cas.cn/dkxzz/HEPS/download/
HEPS-High voltage line station, official website page | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262831.html
HEPS-High voltage line station, person in charge | Li Xiaodong (lixd@ihep.ac.cn)
HEPS - high voltage line station, PILOT contact | Zhang Fei (fzhang@ihep.ac.cn)
HEPS-high voltage line station, line station overview | The ID31 high-pressure beamline station is one of the first special line stations built by the High Energy Synchrotron Radiation Source (HEPS). It is oriented to the needs of in-situ characterization of material structures under extreme conditions and aims to support cutting-edge research in the fields of materials science, earth and planetary science, physics and chemistry, especially to solve major scientific problems related to extremely high pressure (such as the exploration of metallic hydrogen). Relying on the advantages of high energy and low emittance of the HEPS 6 GeV storage ring, the line station can provide high-flux X-ray probes with a wide energy range of 20–75 keV and multi-scale focusing spots (200 nm–20 μm). Combined with three independently configured experimental platforms and the advanced Eiger2
HEPS-High Voltage Line Station, Design Optimization (DAC Demand) | The sample of the DAC experiment is extremely small, and the diamond anvil absorbs seriously in the low energy range. Therefore, obtaining a high signal-to-noise ratio diffraction signal requires small, compact, and high-energy X-rays. In addition, ultra-high-pressure diffraction experiments have extremely high requirements on the stability of the beam line. For this reason, we have optimized the beam line design specifically to meet the needs of DAC users. The use of a horizontally reflective pseudo-groove dual crystal monochromator (CCM) equipped with Si(111) and Si(220) can cover the energy range of 20-75 keV; the use of compound refractive lenses (CRLs) and K-B focusing mirrors can achieve focused spots from micron scale to sub-micron.
HEPS - High Voltage Line Station, General High Voltage Platform (GPT) | Located upstream in the first experimental shed, it provides 20, 30, 40, 50 and 75 keV monochromatic light, and achieves submicron-micron light spots through CRLs. GPT is equipped with a high-precision, large-load-bearing sample positioning system, supports a variety of extreme environments such as resistance heating, cryostat and dynamic loading, and is compatible with large-scale equipment such as Paris-Edinburgh presses. Combined with the Eiger2 detector, it supports <1 kHz dynamic loading experiments (P < 300 GPa, 4–2000 K).
HEPS-High Voltage Line Station, Laser Heating Platform (LHT) | Located downstream in the first experimental shed, it is dedicated to double-sided laser heating experiments with 20, 30, and 40 keV micron focused spots. It has continuous and pulse modes, the temperature can reach more than 4000 K, and is equipped with an accurate temperature measurement system to support materials research in high-temperature-high-pressure conditions.
HEPS-High Voltage Line Station, Submicron Focused Platform (Sub‑μT) | Located in the second experimental shed, it provides 20–40 keV monochromatic light with <200 nm focusing via a pre-shaped K‑B mirror. The experimental station is equipped with a high-precision air-floating turntable and a three-dimensional piezoelectric nano-positioning system, which is suitable for carrying out extremely small-volume samples and imaging experiments.
HEPS-high voltage line station, online measurement and efficiency improvement | All platforms are equipped with online ruby fluorescence and Raman spectroscopy systems for pressure measurement, as well as air film pressurization systems to improve experimental efficiency and flexibility.
HEPS-high voltage line station, detector | The experimental station is equipped with the X-ray detector Eiger XE 16M CdTe.
HEPS-High Voltage Line Station, Development Plan | At the beginning of the design, the line station fully considered the scalability of future development, and reserved space in the layout of the optical and experimental stations to facilitate the development of new experimental methods such as coherent diffraction, emission spectrum, and imaging, ensuring that it maintains international advanced levels in the long term.
HEPS - High Voltage Line Station, Experimental Method - Powder XRD | Powder XRD
HEPS-High Voltage Line Station, Experimental Method-Single Crystal XRD | Single crystal XRD
HEPS-High Voltage Line Station, Experimental Methods-Multiple Particle XRD | Multi-grain XRD
HEPS - High Voltage Line Station, Experimental Method - Radial XRD | Radial XRD
HEPS - High Voltage Line Station, Experimental Methods - High Voltage PDF | High voltage PDF
HEPS - High Voltage Line Station, Experimental Methods - Time-Resolved Diffraction | time resolved diffraction
HEPS-high voltage line station, sample environment-low temperature | low temperature
HEPS - high voltage line station, sample environment - resistance heating | Resistance heating
HEPS-high voltage line station, sample environment-dynamic loading | dynamic loading
HEPS-high voltage line station, sample environment-double-sided infrared laser heating | Double-sided infrared laser heating
HEPS-high voltage line station, sample environment-PE type press | PE type press
HEPS-high voltage line station, main parameters-light source | undulator IVU19.9
HEPS-high voltage line station, main parameters-total length | 4.0 m
HEPS - high voltage line station, main parameters - cycle length | 19.9 mm
HEPS-high voltage line station, main parameters-FWHM opening angle @Kmax | 0.26 (H) × 0.11 (V) mrad^2
HEPS-high voltage line station, main parameters-fundamental light source FWHM size @Kmax | 21.9 (H) × 9.0 (V) μm^2
HEPS-high voltage line station, main parameters-fundamental light source FWHM divergence angle @Kmax | 13.6 (H) × 11.8 (V) μrad^2
HEPS-high voltage line station, main parameters-photon brightness @30 keV | 1.7 × 10^{22} phs/s/mm^2/mrad^2/0.1%B.W.
HEPS-high voltage line station, main parameters-photon flux @30 keV | 3.8 × 10^{14} phs/s/0.1%B.W. @16 μrad × 16 μrad
HEPS-high voltage line station, precautions during experiments | • Ensure that at least one experimenter is online during the entire machine time. • Do not bring any food or drinks into the experimental hall, and do not eat or drink in the experimental hall. • During this machine time, other people who are not included in the "experimental personnel" list on the experimental safety audit form are not allowed to use the station equipment. • Ensure that the experiment ends within the allocated machine time to ensure normal experiments for subsequent users.
HEPS - High Voltage Line Station, Sample Preparation - About Samples | DACs & Holders; Resistive Heating; Low Temperature Systems.
HEPS-High Voltage Line Station, Sample Preparation-Key Parameters of Low Vibration-Free Liquid Helium Closed Cycle Low Temperature and High Pressure Device | The temperature can be reduced to 4 K, and the temperature stability is better than ±100 mK; liquid helium circulation cooling; equipped with a membrane pressurization and pressure relief module, which can remotely control the pressure loading; the stereoscopic angle of the diffraction window is ≥50; compatible with DAC: Piezzo DAC.
HEPS - high voltage line station, supporting laboratory - availability | Users can use the user-assisted laboratory before the experiment (reservation required) and during the experiment.
HEPS-high voltage line station, supporting laboratory-equipment | High-power microscope; visual high-pressure inflation system; glove box; laser drilling equipment; offline ruby pressure measurement system.
HEPS - high voltage line station, data format | ID31 records data in compressed .hdf5 file format, which can be viewed directly with software such as Dioptas, and provides format conversion codes to convert .hdf5 to .tif and .esperanto.
HEPS-High Voltage Line Station, Data Processing Software | Standard software used for data analysis are Dioptas and CrysAlisPro.
HEPS-High Resolution Spectroscopy Line Station, official website | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262830.html
HEPS - High Resolution Spectroscopy Line Station, title | High resolution spectroscopy line station
HEPS - High Resolution Spectroscopy Line Station, alias/abbreviation | Hard X-ray High Energy Resolution Spectroscopy Line Station (H^2O)
HEPS-High Resolution Spectroscopy Line Station, Line Station Introduction | The Hard X-ray High Energy Resolution Spectroscopy Line Station (referred to as H^2O) is my country's first line station focusing on hard X-ray inelastic scattering spectroscopy experiments. It mainly provides nuclear resonance scattering (NRS), X-ray Raman scattering (XRS), resonant inelastic scattering (RIXS) and other spectroscopic methods. Focus on the detection of nuclear energy level hyperfine structure, phonon state density, core energy level electronic transition and valence electron excitation. Mainly serving multi-disciplinary cutting-edge research in quantum science, energy science, materials science, condensed matter physics, chemistry, biochemistry, earth science, high-voltage science, environmental science, etc.
HEPS-high resolution spectroscopy line station, line station location and configuration | The line station is located in HEPS low Beta sector ID33, using vacuum internal undulator light source; the main optical components and experimental station equipment include: white light mirror, liquid nitrogen cooling monochromator (Si(111)/Si(311) switchable), compound refractive lens (CRL), multiple types of high energy resolution monochromator (ΔE/E=10^{-6}~10^{-7}), K‑B focusing mirror, NRS Experimental devices (non-focused/focused and time-resolved detection systems), large solid angle multi-function spectrometer, and a variety of in-situ sample environments (high pressure, high and low temperature, magnetic field, electrochemical workstation, etc.).
HEPS-high resolution spectroscopy line station, energy range keV | 7–26 keV
HEPS-High Resolution Spectroscopy Line Station, Energy Resolution-NRS | 1 meV @14.4 keV
HEPS-High Resolution Spectroscopy Line Station, Energy Resolution-XRS | 0.7 eV @9.7 keV
HEPS-High Resolution Spectroscopy Line Station, Energy Resolution-RIXS | 100 meV @8.9 keV
HEPS - high-resolution spectroscopy line station with the smallest focused spot size | 2 µm × 2 µm
HEPS-High Resolution Spectroscopy Line Station, Luminous Flux at Sample-NRS | ~1.5×10^{10} phs/s/meV @14.4 keV
HEPS-High Resolution Spectroscopy Line Station, Light Flux at Sample-XRS | ~3×10^{13} phs/s @9.7 keV
HEPS-High Resolution Spectroscopy Line Station, Light Flux at Sample-RIXS | ~5×10^{12} phs/s @8.9 keV
HEPS-High Resolution Spectroscopy Line Station, person in charge | Xu Wei (xuw@ihep.ac.cn)
HEPS-high resolution spectroscopy line station, line station pictures | [Picture placeholder]
HEPS-High Resolution Nanoelectronic Structure Line Station, official website | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262839.html
HEPS - High Resolution Nanoelectronic Structure Line Station, Title | High-resolution nanoelectronic structure line station
HEPS-High-resolution nanoelectronic structure line station, line station introduction-High-NESS | High-Resolution Nanoscale Electronic Structure Research Line Station Spectroscopy (High-NESS for short) uses an undulator insert to provide polarization-adjustable high-brightness soft X-rays, covering the At the same time, it covers the K-edge of light elements such as C, N, and O, the L-edge of transition metals, and the M-edge of rare earth elements. It is the best experimental light source for the study of strongly correlated systems, and can provide rich physical information for surface and interface research on various new quantum materials.
HEPS-High Resolution Nanoelectronic Structure Line Station, working mode | It is planned to be designed with two working modes: nano-angle-resolved photoelectron spectroscopy (Nano-ARPES) and micron angle-resolved photoelectron spectroscopy (μARPES) to ensure the different optimization requirements of the light source for different experimental needs of experimental users, and the light spots of the two experimental modes are confocal and can be freely switched without destroying the vacuum.
HEPS-High Resolution Nanoelectronic Structure Line Station, person in charge | Wang Jiaou (wangjo@ihep.ac.cn)
HEPS-optical test line station, official website page | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262846.html
HEPS-optical test line station, line station introduction | HEPS has extremely high performance requirements for optical components and detectors, many of which challenge the limits of current processes and detection. The optical test beamline is a dedicated beamline used for the research and testing of key optical, detection and experimental equipment. It mainly tests focusing equipment, monochromatic equipment, and detectors, and conducts surface shape accuracy, mechanics, thermal and optical performance analysis, phase impact and control, and stability impact and research. ; Carry out verification research on new experimental methods, provide experimental conditions for white light, pink light, monochromatic light and focusing. The energy range can be extended to 300keV, which can meet the needs of various thermal power load experiments and can meet the testing needs of various focusing schemes (such as KB, CRL, etc.).
HEPS-Optical Test Line Station, person in charge | Jia Quanjie (010-88235994 jiaqj@ihep.ac.cn)
HEPS-X-ray absorption spectroscopy line station, official website | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262833.html
HEPS-X-ray absorption spectroscopy line station, title | X-ray absorption spectroscopy line station
HEPS-X-ray absorption spectroscopy line station, introduction to the line station | The X-ray absorption spectroscopy line station is a high-performance hard X-ray beamline led by an insert and optimized for X-ray absorption spectroscopy and related derivative experimental methods. The basic goal of the design is to have strong versatility, large energy coverage, and a high-throughput, high-brightness XAFS experimental system to provide users with a reliable, stable, and easy-to-operate absorption spectroscopy experimental platform. The platform has conventional experimental methods for X-ray absorption spectroscopy, has time resolution and spatial resolution capabilities, and can detect micro-area/surface/interface structures. At the same time, the online station also provides users with a variety of in-situ sample environments, and combined with experimental methods such as XRF/XRD/FTIR/mass spectrometry, it can provide richer sample information. At the same time, the design provides the possibility for the development of other experimental methods and new experimental methods. The target energy covers 4.8–45 keV (covering the K-edge XAFS energy range of element 22, Ti-60, Nd, and covering the L-edge of element 55, Cs and above).
HEPS-X-ray absorption spectroscopy line station, main indicator-energy range | 4.8–45 keV
HEPS-X-ray absorption spectroscopy line station, main indicator-energy resolution | 2×10^{-4}(Si(111)); 4×10^{-5}(Si(311))
HEPS-X-ray absorption spectroscopy line station, main indicator-spot size | ~2 mm × 2 mm; ≤0.35 μm × 0.35 μm (focus mode)
HEPS-X-ray absorption spectroscopy line station, main indicator-luminous flux | 2×10^{13} ph/s @10 keV; 5×10^{12} ph/s @10 keV (focus mode)
HEPS-X-ray absorption spectroscopy line station, person in charge | Zheng Lirong (zhenglr@ihep.ac.cn)
HEPS-General Environmental Spectroscopy Line Station, official website page | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262840.html
HEPS - Universal Environmental Spectroscopy Line Station, Title | Universal Environmental Spectroscopy Line Station
HEPS-Universal Environmental Spectroscopy Line Station, Person in Charge | Zheng Lei (zhenglei@ihep.ac.cn)
HEPS - Universal Environmental Spectroscopy Line Station, PILOT Contact | Ma Chenyan (macny@ihep.ac.cn)
HEPS-Universal Environmental Spectroscopy Line Station, Science Objective | Stable and efficient absorption spectrum testing in the X-ray energy range of 2.1 keV ~ 10 keV to obtain information including the chemical form, geometric structure and electronic structure of substances.
HEPS - Universal Environmental Spectroscopy Line Station, main features | Covering X-rays in the medium energy range, it can develop the K absorption edge of low Z elements (S, P, Cl, Ar, K, Ca) and heavy metal elements (Cu, Fe, Co, Ni, Mn, etc.), as well as the L and M absorption edges of transition metal elements; the sample forms are diverse (solid, liquid, gas, mixed state); the sample environment is diverse: in-situ electrochemistry, electrocatalysis, high and low temperature, etc.
HEPS - Universal Environmental Spectroscopy Line Station, main parameters - energy range (keV) | 2.1–10 keV
HEPS - Universal Environmental Spectroscopy Line Station, main parameter - energy resolution (ΔE/E) | ~2E-4
HEPS-General Environmental Spectroscopy Line Station, main parameter-minimum spot (H×V) | 20 µm × 20 µm
HEPS-General Environmental Spectroscopy Line Station, main parameter-maximum spot (H×V) | 15000 µm × 1500 µm
HEPS - Universal Environmental Spectroscopy Line Station, main parameter - luminous flux | 1E10 ~ 1E12 phs/s
HEPS - Universal Environmental Spectroscopy Line Station, main parameters - sample environment | high temperature/low temperature
HEPS - Universal Environmental Spectroscopy Line Station, Main Parameters - Detector | 7 yuan SDD, Lytle
HEPS-Universal Environmental Spectroscopy Line Station, Main Experimental Methods | XAFS - X-ray absorption fine structure
HEPS-Universal Environmental Spectroscopy Line Station, Main Experimental Methods | XANES - X-ray absorption near-edge structure
HEPS-Universal Environmental Spectroscopy Line Station, Main Experimental Methods | XRF - X-ray fluorescence
HEPS-universal environmental spectroscopy line station, line station equipment-experimental station composition | The experimental station consists of two parts: the vacuum chamber and the EXAFS device.
HEPS - Universal Environmental Spectroscopy Line Station, Line Station Equipment - Vacuum Chamber | The vacuum degree can reach below 1E-4 Pa. It is generally used for absorption edge testing of elements in the low energy range of 2.1-5.0 keV. The modes include the total electron yield method (TEY) and the partial fluorescence method (PFY). Vacuum and helium atmospheres can be converted into each other to reduce air absorption. Liquid samples of low-energy elements can be tested in a helium atmosphere.
HEPS-Universal Environmental Spectroscopy Line Station, Line Station Equipment-EXAFS Measurement System | In the atmospheric environment, it is generally used for absorption edge testing of elements higher than 5 keV. Transmission method (TM) and fluorescence yield method (PFY) can be used; it has three ionization chambers, and the test sample is placed between the front ionization chamber 1 and the rear ionization chamber 2. A standard filter disk is placed in front of the rear ionization chamber 3 to calibrate the energy in real time.
HEPS-universal environmental spectroscopy line station, line station equipment-sampling robot | The EXAFS device is equipped with a 6-axis sampling manipulator and sample tray, which can accommodate 50 samples at the same time, and the sample change time is <10 seconds; the control system is embedded in the spectrum acquisition program, which can automatically replace samples and continuously collect spectra without manual intervention.
HEPS-Universal Environmental Spectroscopy Line Station, Line Station Equipment-KB Focusing Lens | A KB focusing mirror is installed behind the experimental station to focus the light spot to about 13 µm × 14 µm, with a working distance of about 200 mm, which is suitable for small focus spot experimental tests.
HEPS-Universal Environmental Spectroscopy Line Station, Sample Environment-High Temperature Gas-Solid Phase In-Situ Device | It is used for solid samples in the atmospheric environment. It can be heated from room temperature to 873 K with an accuracy of ±1°C. The temperature can be controlled locally or remotely. It is compatible with transmission and fluorescence modes. A gas circulation device is installed on the top to pass in reaction gases to achieve in-situ solid-gas phase reaction. A sapphire window installed on the same side of the sample surface can introduce external trigger light to achieve photocatalytic reaction conditions.
HEPS - Universal Environmental Spectroscopy Line Station, Sample Environment - Cryostat | Liquid nitrogen refrigeration, temperature control range 80–300 K (room temperature), liquid nitrogen capacity 1.6 L; target temperature can be reached in 30 minutes at the fastest, temperature control stability ±0.1 K; equipped with two sample rods (can hold 2 samples each time), easy to replace in about 5 minutes; good sealing, stable at low temperatures for more than 10 hours; inner window diameter 15 mm (120 µm high-purity Be film window), outer window diameter 35 mm (120 µm Mylar window); equipped with fluorescence and transmission sample holders.
HEPS - Universal Environmental Spectroscopy Line Station, Detector - Configuration | Two 7-yuan SDD detectors are located in the vacuum chamber and the EXAFS device respectively; equipped with a large-area fluorescence detector suitable for fluorescence signal collection of low-concentration samples.
HEPS-Universal environmental spectroscopy line station, sample preparation equipment | Mortar, tape, automatic tablet press.
HEPS - Universal Environmental Spectroscopy Line Station, Data Acquisition System | The self-developed absorption spectroscopy data acquisition system enables free switching of dual acquisition spectroscopy systems; it is embedded with functions such as automatic gas distribution and automatic sample change, making it user-friendly.
HEPS - Universal Environmental Spectroscopy Line Station, Data Processing Software | Two sets of absorption spectrum data analysis software developed by HEPS Computing and Network Communication System support real-time online or offline processing of XAS match, PCA‑LCF and extended edge fitting.
HEPS-Universal Environmental Spectroscopy Line Station, Examples of Application Areas-In-situ Electrochemistry Research on Solid-State Li-S Batteries | All-solid-state lithium-sulfur batteries (ASSLSB) have the potential of high specific energy, high safety and low cost. Pang Quanquan's research group at Peking University designed a composite solid-state sulfur cathode to trigger a rapid sulfur solidification reaction, allowing the battery to exhibit excellent fast-charging performance. The K-edge XANES of S was used to test the valence state evolution of the sulfur cathode during the charge and discharge process: S in the original electrode was converted into Li2S and Li2Sx after the first discharge; after full charge, Li2S and Li2Sx were completely converted into sulfur; after another discharge, the sulfur was converted into Li2S and Li2Sx, showing good reversibility. Relevant results were published in Nature (2025), DOI:10.1038/s41586-024-08298-9.
HEPS-General Environmental Spectroscopy Line Station, Examples of Application Fields-Research on the Reduction Mechanism of Mercury Pollution | The concentration and isotopic characteristics of Hg2+ and Hg0 in the peat soil at different depths in the peatland system in northern Sweden were studied to reveal the redox transformation process of mercury after deposition: dry and wet sedimentation inputs 36% of mercury into peatlands, and peat vegetation absorbs 64%; surface photochemical reduction causes about 30% of Hg2+ to be reduced to Hg0 and released to the atmosphere every year; dark oxidation of organic matter above the water table reduces Hg0 Concentration; supersaturated Hg0 below water table More from direct photochemical reduction of Hg2+ in rainwater. Related research used the BSRF 4B7A line station for sulfur speciation analysis and was published in Nature Communications (2023) 14:7389.
HEPS-auxiliary laboratory, official website page | https://ihep.cas.cn/dkxzz/HEPS/zzjs/gsxz/201903/t20190328_5262847.html
HEPS - Auxiliary Laboratory, Introduction | Regarding the content and functions of the experimental stations of the first batch of beam lines, the auxiliary laboratories will mainly include five aspects: biology, materials, chemistry, environment and in-situ sample environment support laboratories. The main functions of the laboratory are user support, sample preparation and temporary storage for synchrotron radiation experiments. Each laboratory has a strict laboratory safety system, with dedicated personnel responsible for systematically managing the use of the laboratory, sample preparation and storage, etc. At the same time, the laboratory is equipped with radiation monitoring equipment and safety interlock alarm systems.
HEPS-Universal Environmental Spectroscopy Line Station, Person in Charge | Zhang Jing (010-88235980 jzhang@ihep.ac.cn)